# DRAFT ONLY

# REVISION IN PROGRESS

# INFANT HEARING PROGRAM

# AUDIOLOGIC ASSESSMENT PROTOCOL AND GUIDELINES 2005

Contact:  	Martyn Hyde, PhD

Otologic Function Unit

Mount Sinai Hospital, Toronto

416.586.4510

Contents

Executive Summary  ……………………………………………………………….…... 3

Assessment Protocol and Guidelines

Background	……………………………………………………………………….………. 6

ABR Testing – General Considerations	………………………………….……… 11

Threshold ABR with Tonepips

Click ABR Measurement

Measurement of RECD

VRA Testing – General Considerations

VRA Test Protocol

DPOAE

Middle Ear Analysis

Overall Diagnostic Inference

Records and Reports

Suggested Readings

Appendix A – IHP High Risk indicators

Appendix B – ABR Technical Parameters

Appendix C -  VRA Technical Summary

Appendix D – Calibration and Protocol Files

Appendix E – RECD Predictions

Appendix F – RECD Measurement Audioscan RM500

SUMMARY

This protocol addresses audiologic assessment ('Assessment') of infants registered in the Ontario Infant Hearing Program (IHP).

The purpose of the Assessment is to determine the presence of the target ‘Permanent Childhood Hearing Impairment’ (PCHI) and to provide an adequate audiometric basis for informed family choice of effective options for communication development before six months of age, wherever feasible.

The nominal target PCHI is currently >25 dB of permanent hearing impairment at any frequency in the range 0.5-4 kHz, in either ear.  The IHP definition of PCHI includes conductive impairment associated with structural anomalies of the ear but does not include impairment attributable to otitis media.

The specific objectives of the Assessment are to obtain valid and accurate estimates of ear-specific, frequency-specific hearing thresholds, and to determine the nature of any hearing impairment present (conductive, sensorineural, mixed).

Assessments shall be conducted by audiologists registered with CASLPO who are also registered with the IHP as having received approved training in this Assessment protocol.

Assessments shall be conducted only using audiometric equipment approved by the IHP, configured, calibrated and maintained according to IHP guidelines, and using operating supplies approved by the IHP.

Assessments shall be conducted in a manner that reflects the IHP core values of fully informed family/caregiver choice and consent, timely provision of unbiased information based on scientific evidence, and cultural sensitivity.

Assessments shall comply with all current guidelines and protocols issued by CASLPO and by the IHP.  Updates to this protocol will be issued as required, and the 'current' protocol includes all updates distributed at least one month before the date of Assessment.  Should this protocol and any CASLPO professional practice guideline be perceived as in conflict, the CASLPO guideline shall apply and IHP management should be advised of the conflict.

Departures from this protocol may be appropriate in individual infants and under special circumstances, and their nature and rationale shall be documented.  The IHP may seek to review documentation and clinical records involving such departures from this protocol.

ABR Based Assessments

Most infants will first present for Assessment at a corrected age of three months or less and will have failed click AABR screening.  Some will present later, through repeat screening or referral of infants at risk for hearing impairment.

Wherever feasible and not medically contra-indicated, Assessments should be targeted to occur at a corrected age of about 8-12 weeks, and in the case of NICU graduates after extended hospital stays, to occur at about 2-4 weeks after discharge from hospital.

VRA Based Assessment

Where developmentally appropriate, visual reinforcement audiometry (VRA) shall be used to obtain behavioural estimates of hearing sensitivity.  Wherever feasible, VRA shall address frequency-specific and ear-specific thresholds by air conduction, and also by bone conduction if indicated.

Within the IHP, an infant previously identified as having a permanent childhood hearing impairment (PCHI) on the basis of the ABR-based assessment protocol would require behavioural audiometric assessment prior to reaching 12 months of age.  Infants of at least 6 months corrected age and up to approximately 24-30 months of age (or until the child can reliably perform play audiometry) are candidates for VRA.

Wherever feasible, Assessments shall include at least ALL of the following, in each ear:

Tonepip ABR thresholds at 0.5 kHz and 2 kHz by air conduction (AC) and also by bone conduction (BC), or VRA based behavioural thresholds at 0.5 kHz and 2 kHz by air conduction (AC) and also by bone conduction (BC), where indicated by this protocol;

DPOAE amplitude and noise floor measurements at 1kHz through 4 kHz;

Tympanometry using probe frequencies of 226Hz or 1kHz, as is appropriate by age;

Ipsilateral middle-ear muscle reflexes at 1 kHz using a probe frequency of 1kHz.

Wherever feasible, ABR, VRA, and OAE testing shall be done in an audiometric soundroom that satisfies current ANSI standards for manual pure tone audiometry.  Testing in any other environment will not qualify for IHP funding unless specifically approved by the IHP.

ABR testing shall be attempted during natural sleep, unless medically contraindicated.  Failure to achieve satisfactory measurement conditions in natural sleep may lead to a recommendation for Assessment under sedation or general anesthesia.

Assessments under sedation shall comply with generally accepted standards of care and all local safety protocols.  The IHP recommends informed consent, medical referral and specification of sedative and dosage, administration by medical/nursing staff, appropriate supervision of the child post-medication and adequate access to emergency services. Wherever feasible, air conduction (AC) tonepip ABR thresholds shall be obtained using insert earphones.  Supra-aural earphones shall be used only if insert earphones are specifically contraindicated.

The required threshold at any frequency is 25 dB Estimated Hearing Level (EHL, see text for definition). [The required minimum response level by VRA is 30dBHL]

If the AC or BC ABR thresholds at 0.5 kHz and 2 kHz differ by 25 dB or more in any ear, threshold measurement at 1 kHz shall be attempted in that ear.

If an AC threshold at 2 kHz is 40 dB EHL or greater in any ear, threshold measurement at 4 kHz shall be attempted in that ear.

If any AC threshold at 0.5 kHz or 2 kHz is elevated by more than 10 dB above the required minimum stimulus levels specified here, then the corresponding BC threshold determination shall be attempted.

BC testing shall be done with superoposterior placement and securing of the transducer above the individual test ear. The transducer shall not be hand-held during the test. For ABR, two-channel EEG recording (forehead to ipsilateral and contralateral mastoid) will be used.

If there is no detectable ABR to tonepips at the highest available intensity levels at all frequencies measured in any ear, by AC and BC (if indicated), then an otoneurologic ABR test with high-intensity AC click stimuli shall be attempted in that ear.

Otoneurologic ABR testing shall include separate records with condensation and rarefaction clicks, to permit evaluation of the cochlear microphonic potential (CM).  It shall also include procedures to distinguish CM from stimulus artifact.

Normal hearing in any ear is defined in the IHP context by reliable response detection at 30 dB EHL/HL at 0.5 kHz and 2 kHz, and at any other frequency if tested.

One or more reliable thresholds greater than 25dB EHL at any frequency in the range 0.5 – 4 kHz in either ear are sufficient to define an infant as having a hearing impairment, and are necessary but not sufficient to define an infant as having PCHI.

The audiologist shall fully inform the family of the nature and implications of test results and shall provide all relevant IHP informational materials.  Contact information for the relevant IHP family support worker shall be provided to families of infants with confirmed PCHI.

The audiologist shall complete the appropriate IHP Assessment report form and send it to the local coordinating PSL agency in a timely manner.

All audiological records shall be maintained in a manner satisfying the requirements of both CASLPO and the IHP.  The Assessment records shall be maintained in both hardcopy and electronic formats and records shall be sufficient to fully specify the subject, tester, test date and location, and the test parameters.

Duplicate Assessments under IHP funding and intended to provide a second opinion are discouraged, but consultation for clinical decision support within the IHP is encouraged.  Exceptional cases will be evaluated by IHP management, on an individual basis.

Communication development procedures conducted under IHP funding shall be based on audiometric data obtained in full compliance with this Assessment protocol.

AUDIOLOGIC ASSESSMENT PROTOCOL AND GUIDELINES

Background

This document addresses audiologic assessment of infants registered with the Infant Hearing Program (IHP).  The contents are based on: (i) a number of workshops for all IHP audiologists from Dec 2000 to March 2004 (ii) systematic reviews of scientific and clinical literature, (iii) consultations with experts worldwide, and (iv) extensive experience with tonepip ABR and other pediatric audiometric procedures, in Ontario.

To justify the resource expenditure required by a universal program, there is a need to achieve a high quality of care and consistency of practices across the province.  Procedures defined here as 'required' specify a standard of care for IHP Assessments.  Some procedures are recommended but not required, and in other areas there may be insufficient evidence even for a recommendation.  The IHP acknowledges that individual infants and special circumstances of testing may require clinical judgement and adaptations of standard procedures.  When the standard of care is not followed, a documented rationale for the departures may be required by the IHP.

The core values of the IHP are that service provision should be family-centered, with fully informed family/caregiver (hereafter ‘family’) choice based on unbiased information.  'Family-centered' means that the family’s choices are paramount, not the opinions of the service provider.  The family’s culture must be accorded respect.  The family must be a full partner in the development of an individualized pattern of required services, and must be assisted in making choices among service options on the basis of information that is timely, comprehensible, relevant, complete, unbiased and based on scientific evidence. Interactions among service providers and families must reflect these core values and also must be consistent with information for families and professionals that is provided by the program.

During their path through the program, families will be given information about program options, procedures and possible outcomes.  This will be made available in many of the languages most frequently represented in Ontario.  Families shall be encouraged to consider the evidence carefully in arriving at their choices. In materials supplied by the IHP, areas in which there is a lack of sound, scientific evidence will be identified.  Standard, published methodologies will be used in determining the quality of scientific evidence.

As well as defining the clinical assessment protocol and standard of care, this document outlines the rationale for specific procedures and decision rules, and also provides some information about the operational context and processes.  The clinical protocol itself is based on current evidence about the effectiveness and efficiency of specific procedures.  Thus, it is not static but will evolve and will be improved over time.  In some areas, current evidence is incomplete and interim decisions have been made.  Also, the IHP will continue to evaluate its operations and outcomes, as well as continue to assess new clinical technologies and published scientific data.  Revisions or addenda to this document will be issued as required.

Target Population

Candidates for Assessment will normally be resident in Ontario and will arise by one of three mechanisms: universal newborn hearing screening (UNHS), surveillance of high-risk infants, and referral. The UNHS path will dominate the caseload.  Newborns with risk indicators for hearing impairment will be screened by automated ABR (AABR), ideally just before hospital discharge.  Newborns not known to be at risk will be screened by automated DPOAE (ADPOAE) immediately followed by AABR in ADPOAE referrals. Wherever feasible, this will be done before discharge from the birth hospital admission. Babies not at risk who have a refer result on their first (usually pre-discharge) AABR will be re-screened by AABR, ideally after an interval of about 2-4 weeks.  Infants with a ‘refer’ result in any ear at the AABR screen (at-risk babies) or re-screen (babies not at risk) will attend for Assessment, which ideally will be initiated at about 8-12weeks corrected age.  For NICU graduates who have had extended stays in hospital, Assessment should be initiated within four weeks of discharge from hospital, wherever feasible.

Some infants who pass the newborn screening will be at risk for non-congenital hearing impairment according to perinatal indicators (Appendix A).  These infants will be recalled for surveillance assessments (See Audiology Surveillance Protocol), until they are 18 months of age.  Some infants who have early-onset or progressive impairment will be identified by this re-screening.

The third path to assessment is physician referral into the IHP of children with post-atally acquired or expressed risk indicators (Appendix A).  The IHP will provide information to both families and physicians in order to improve identification of such infants.  Children detected by this route may be up to 24 months of age at the time of physician referral; eligibility for referral into the IHP ceases at the second birthday.  Referral must be based on evidence of a valid risk indicator.  Otitis media (OM) is NOT an acceptable solitary indicator.  An evaluation of possible procedures in relation to OM and IHP candidacy is ongoing.

Because of resource constraints, at the present time children who are aged two or more at the time of identified service need will receive audiology services through existing systems and programs other than the IHP.

# Target disorder

The nominal target disorder for the IHP is ‘permanent’ childhood hearing impairment (PCHI) in either ear of greater than 25 dBHL at any frequency in the range 0.5 to 4.0 kHz.  A hearing impairment is considered ‘permanent’ by the IHP if it is sensorineural or mixed, or has ‘structural’ conductive etiology (such as atresia or middle ear malformation).  It is considered appropriate to identify children with unilateral PCHI because: (i) they are at risk for bilateral PCHI, (ii) they are at risk for significant disability should the normal ear acquire a conductive disorder, and (iii) specific strategies are indicated to enhance communication development in children with unilateral impairment.

In practice, the ‘effective’ target disorder for a UNHS program is dictated by the operating characteristics of the screening tests used.  AABR screening is currently done using clicks. The click level can be selected by IHP management and currently is 35 dBnHL.  However, the click ABR will reflect the best puretone sensitivity in the frequency range 1-4 kHz, so children with hearing losses at specific frequencies may be missed.  Similarly, DPOAE screening typically addresses frequencies of 1.5 kHz or 2 kHz and higher, so hearing losses at 1 kHz and below will be missed.  Current IHP settings for DPOAE screening are 1.5, 2, 3 and 4 kHz with a 3 out of four pass rule.

Purpose of Assessment

The objectives of the audiologic Assessment are

- To determine the presence of the target PCHI.

- To estimate puretone thresholds in each ear.

- To determine the type of hearing loss (audiologic diagnosis*).

- To determine audiologic candidacy for non-medical communication development options, such as amplification.

- To assess referral needs (medical and non-medical).

- To facilitate family access to required supports, information and service options for communication development.

- The term 'diagnosis' is used here not in the sense of an etiologic diagnosis, which may only be communicated by a physician under the RHPA, but in relation to the assignment of categorical labels to conventional hearing loss descriptors, such as 'conductive', 'sensorineural', and 'mixed'.

# General Approach to Assessment

IHP Assessments must be conducted by an audiologist who is registered with the College of Audiologists and Speech-Language Pathologists of Ontario (CASLPO) and who is registered with the IHP as having received specific, required training in this protocol. Should this protocol and any professional practice guideline adopted by CASLPO be in conflict, the CASLPO guideline shall apply.  Where this protocol is more specific, or where no CASLPO guideline exists, this protocol shall apply.

Fees for IHP Assessments will be administered by the regional IHP coordinating agency, according to a schedule subject to program per-contact maxima and agreed and contracted locally.  Procedures conducted in compliance with these guidelines will normally qualify for IHP reimbursement, and shall not concurrently be billed to OHIP or any other provider.

In this protocol, the emphasis is upon techniques for Assessment in the first two years of life.  In the first six months,, Assessment is based on objective, physiologic measures, especially tonepip ABR.  It is usually possible to obtain accurate, frequency-specific and ear-specific puretone threshold estimates by these techniques, even in neonates.  In most cases, tonepip ABR methods can provide audiometry that is fully sufficient to consider proceeding with communication development options, such as amplification.  If proper ABR techniques are used, then unless there is a specific indication of possible unreliability of ABR findings, it is inappropriate to defer initiation of communication development options pending ‘behavioural confirmation’ of threshold estimates.

Throughout the first year of life, the likelihood of successful ABR testing in natural sleep decreases, and over about six months corrected age it can be quite difficult to achieve.  Testing under sedation is usually an option, although in some parts of the province significant travel may be required to access such services.   Also at about six to eight months corrected age, in children with normal cognitive development it may become feasible to assess hearing behaviourally.  Initially, such testing is limited in reliability and scope, but it is possible to obtain useful estimates of frequency-specific and ear-specific response thresholds by air and bone conduction in children under one year of age, given normal development.

As behavioural testing becomes increasingly accurate and specific, there should be a gradual transition from reliance upon electrophysiologic measures in most infants to reliance on behavioural measures in most infants.  For those infants with significant neurologic, cognitive or behavioural disorders, however, electrophysiologic methods may remain the primary audiometric tool.

While current clinical, behavioural audiometric methods are generally inaccurate in the first six months of life, they may be useful qualitatively in addressing the beliefs of the family about the child’s hearing.  Family observations are unreliable as indicators of hearing sensitivity at that age, but they must be considered carefully and managed tactfully and informatively.  In older infants, family reports of behaviours related to hearing may provide important, even crucial, information.

It is necessary to obtain and weigh all types of information that may assist valid and accurate audiologic assessment.  Errors in Assessment can have very serious consequences and every possible effort must be made to avoid them.  Major errors that have been reported in infant assessments include misinterpretation of ABR waveforms (both false-positive and false-negative response detection), and misinterpretation of VRA results, such as underestimation of hearing loss arising from false-positive responses. Accordingly, total reliance on a single, uncorroborated measure of hearing must be avoided wherever possible.  A multi-component approach to Assessment is required, so that redundancy of information can provide cross-checks and reduce the likelihood of major errors.  Discrepancies among test results must be addressed by critical review and by further testing, wherever feasible.

It is also important to note that audiologic Assessment of an infant with PCHI is not an event, but a process.  Even if complete and apparently accurate audiometry were obtained at three months, occasional follow-up audiometry is usually appropriate to confirm the early measurements, to refine threshold estimates and to detect and quantify possible changes in hearing.  Little is known about patterns of progression of PCHI in infancy, and such information is important not only with respect to the individual child but also for ongoing optimization of services.

Wherever feasible and indicated, complete Assessment of the young infant should include:

- Physiologically-based estimation of frequency-specific and ear-specific hearing thresholds.  Currently, this shall be done using tonepip ABR methods, including bone-conduction measurements,

- Objective measurement of cochlear function, currently including DPOAEs as well as cochlear microphonic (CM) potentials,

- Objective middle-ear analysis (MEA), currently including acoustic immittance measurement (tympanometry, with low- and high-frequency probe tones), and middle-ear muscle reflexes (MEMR, ipsilateral, with a high-frequency probe tone),

- Objective measurement of the functional status of auditory brainstem pathway, currently including otoneurologic click ABR, and

- Developmentally appropriate measurement of behavioural response to sound, preferably frequency-specific and ear-specific, by VRA wherever applicable.

The Assessment protocol takes account of the epidemiologic and clinical characteristics of the presenting population.  For example, data from screening programs suggest that many infants who have failed AABR screening will be determined at the Assessment not to have the target PCHI. This is commonly due to a middle-ear disorder that has resolved since the screening referral, or to inadequate screening procedures.  At the Assessment, other infants are likely to have minor, conductive losses, possibly overlayed on the target PCHI.  Accordingly, the protocol is structured to be efficient in such circumstances.

A general theme underlying the clinical strategies identified here is to constantly review the specific clinical information that is most important at any point throughout the course of a clinical assessment, and to implement the precise procedural step that will yield that information in a valid, accurate and efficient manner.  This principle applies to the strategic selection and sequencing of tests, as well as to the detailed tactics within individual tests.

The results of the Assessment are to be summarized on forms provided by the IHP, and transmitted to the regional IHP coordinating agency according to locally-agreed procedures.  In the situation of ongoing Assessments requiring more than one visit, information gained over several visits may be integrated into a single report.  The audiologist is also required to maintain local, clinical records in a manner that will permit detailed analysis in the event of consultative review or quality audit.

# ABR Testing – General Considerations

## Order of Tests

ABR testing is the required core of the Assessment for young infants, but OAE and MEA components are strongly recommended, so the first strategic question relates to test order.  This is a matter for local discretion, and the best order may vary across infants.  Points in favour of OAE testing first, if behavioural conditions permit, are that (i) it may be difficult to obtain a successful OAE test after the child has woken up from the ABR, (ii) having the OAE result immediately informs the tester about possible auditory dys-synchrony, should an absent or abnormal ABR be seen, and this may alter the ABR testing, and (iii) the OAE attempt may remove canal debris and/or improve canal patency.

It can also be argued that MEA results can inform the ABR test as well as improve canal status.  There is a counter-argument that ABR is the core procedure and that doing these other tests up front may irritate the child and prejudice ABR success.  This may be an issue for older infants or those who are inherently irritable.

## Test Environment

The environment for ABR and DPOAE testing should be safe, quiet, of sufficient size (at least 60 square feet is suggested), and adequately ventilated, with low lighting and few visual distractions.  It is normally required that tonepip ABR threshold testing be conducted in an audiometric soundroom satisfying standard ANSI criteria for puretone audiometry.  It is strongly recommended that DPOAE testing also be conducted in the soundroom.  In special circumstances, exceptions may be made, subject to IHP review and acceptance of the proposed environment. In general, Assessments done in any test environment other than an audiometric soundroom will not qualify for IHP funding unless the environment has been previously approved by IHP management.

For alternative test environments to be considered acceptable, conventional puretone threshold audiometry by air conduction using a supra-aural (TDH) earphone in at least five adult subjects with normal hearing should achieve reliable thresholds of 20 dBHL or better at 0.5 through 4 kHz in the precise test situation being proposed.  The environmental review may also include octave band sound level measurements, at the discretion of IHP management. Satisfactory test conditions may be achievable in a quiet, untreated room with little noise from traffic or air handling systems.  If there is a problem, it is most likely to occur at 0.5 kHz.  In ABR and OAE testing, well-fitted insert earphones can provide some protection against environmental noise, but the amount of protection is variable, small and not known exactly.  Because the ABR is an averaged phenomenon, transient sounds may not have a significant effect on the accuracy of threshold estimates.  Furthermore, because the IHP protocol does not pursue ABR thresholds to very low levels, the likelihood of an environmental problem due to steady-state noise is lessened. However, DPOAE measurements are easily compromised by significant room noise, especially at frequencies below 2 kHz.  MEA measurements are not affected by moderate levels of room noise.

Tester and Equipment

The safety and comfort of the infant are the paramount concern, and all reasonable steps shall be taken to ensure them. The infant shall be supervised closely throughout the testing, by an individual who is familiar with all pertinent safety procedures and who has adequate training in handling young infants.  Local protocols must comply with all relevant local safety standards and with generally accepted standards of care.

In many test situations, it is feasible for a single audiologist with appropriate training to conduct tonepip ABR testing.  Where feasible, a prudent arrangement is that the tester and instrumentation are inside the soundroom, together with the infant, who may be in a bassinet, crib or pram.  This may be reassuring to the parents, who may be reluctant to leave the child alone in what they may see as an intimidating test environment.

The current diagnostic ABR/OAE equipment (Biologic NavPro) is laptop-based and noise levels from that unit are not a significant concern.  Printing of records (laser or inkjet) can be done off-line, if necessary, so it is not absolutely necessary to power up the printer during testing.  Alternatively, the printer may be located outside the soundroom, given adequate cable routing though the trap or connector panels.

Attendance of family members/caregivers during ABR testing is a matter of local preference, at the discretion of the IHP audiologist.  In specific situations, the presence of a family member or third party may be desirable for reasons unrelated to test quality, such as to secure compliance or to manage perceived medico-legal risk.  However, family members differ widely in their knowledge and skills related to infants' sleeping habits.  They may distract or excite the child, use inappropriate or even unsafe strategies or manoeuvres, or otherwise compromise efficient and accurate testing.  However, allowing parent(s) into the test room, at least initially, can alleviate their anxiety.  Also, some parents are genuinely skilful, so allowing them to try and get their child to sleep may be helpful.  Thus, a practice of allowing a limited time for parents to encourage sleep, and then leaving the soundroom, is recommended.  This general approach has been successful for many years in several centres.  Primary factors in the choice of approach are the skills of the audiologist at sleep promotion and the pressure of time in local scheduling practices.

## Stimulus Transducers

In the absence of specific contraindications, insert earphones (ER-3A) are the required method of initial choice for ABR testing by air conduction.  Insert phones have several advantages over supra-aural earphones, including reduced stimulus artifact, decreased background noise, less acoustic cross-over, decreased likelihood of collapsed canals, and increased comfort.  Also, it is easier to switch the stimulated ear with inserts, because it is often feasible to test with inserts in both ears.  This is straightforward if the infant is laying supine, which is often the preferred position for very young infants.  More generally, if the inserts are properly placed and held in place by tape, they have little impact on physical positioning of the infant.  Supra-aural earphones are more bulky, are more restrictive in terms of infant position, and require more skill and attention to maintain proper placement.  It can be difficult for the audiologist to apply and maintain one or more supra-aural earphones without assistance, whereas with insert earphones a single tester is usually practicable.  Experience also suggests that insertion and removal of an insert earphone can improve canal patency and/or displace and remove meatal debris or cerumen.

Supra-aural earphones (TDH/MX41) are to be used when insert phones are contraindicated, such as when the ear canals are very small or stenotic or when the infant does not tolerate insertion well.  Careful attention to accurate placement of a TDH earphone is especially important, to ensure appropriate stimulus levels.

Tonepip ABR threshold estimates by bone conduction (BC) are normally required to quantify conductive and sensorineural hearing loss components. In contrast to adult testing, in the infant, intra-cranial transmission losses are sufficiently large that each ear MUST be tested individually, that is, it cannot be assumed that a given mastoid placement stimulates both cochleae equally.

Accurate BC ABR tests require proper placement (supero-posterior to the meatal opening) and stable retention of the transducer, adequate contact force and correct calibration.  To achieve proper force and stability of the bone oscillator, a band of elastic fabric with Velcro attachments is recommended.  Such bands are not available commercially but are easy to construct.  Examples are provided as part of the Assessment training.  The width of material should be sufficient to envelop the transducer and hold it securely in place.  The transducer should NOT be hand-held.  The need for quantitative measurements to ensure application force of about 400 g, as recommended by Yang & Stuart (2000), is currently being evaluated.  At present, application force measurements are optional, provided that the band is applied correctly and with adequate tension.

# Pre-test Preparation

The infant’s behavioural state on arrival for Assessment is important for successful testing in natural sleep.  The family should be made fully aware of the importance of appropriate preparation for testing, and should be given detailed instructions on what to do and what not to do.  Written instructions and telephone confirmation are recommended.  The importance of preparation increases with age, as the infant’s amount of daytime sleep decreases.  Wherever possible, the infant should arrive at the test tired and hungry.  It is normally appropriate to deny sleep and food for at least an hour before testing, where not medically contraindicated.  If the child is being brought to the test by car, it is important that every reasonable effort be made (consistent with safety) to keep the child awake on the journey.  Because of the soporific effect of car journeys on infants, another person than the driver is usually necessary.

On arrival at the Assessment, it is recommended that cursory otoscopy be done, the electrodes be attached and then the infant be fed, before attempting to induce sleep.  MEA testing may be practicable shortly after feeding.  OAE testing may be attempted before the ABR, at the audiologist’s discretion.

## Otoscopy and Cerumen/Debris

The ear canals of young infants frequently contain varying amounts of debris or cerumen.  Hearing testing remains viable unless the canal is completely occluded acoustically, and total occlusion is difficult to determine visually.  Another issue is that the canal may partially collapse, if the infant has been sleeping on one side.  Detailed otoscopy and TM visualization can be difficult in the young infant and are the domain of the experienced physician, but it is recommended that the audiologist conduct at least a cursory otoscopic examination at the outset of the Assessment.  If the canal is obviously occluded, which is infrequent, then removal of the occlusion by an experienced physician is desirable.  However, prompt access to physicians with appropriate skills may be limited.  Usually, successful testing can be accomplished in the presence of substantial debris and/or cerumen, or even given apparently total occlusion.  The decision to undertake testing with insert phones under such conditions is at the discretion of the audiologist.  Supra-aural earphones are an option, with the caveats noted earlier.  It is also possible to undertake bone-conduction testing, but a return visit for AC testing would usually be required to demonstrate normal hearing.

Partial occlusion is not uncommon, but usually does not interfere with accurate testing.  If the insert becomes blocked, it may be found that probe removal withdraws the debris/cerumen and that reinsertion lead to a successful test.  As noted earlier, it can be argued that if the behavioural state of the infant is appropriate, initial MEA and/or DPOAE testing is a convenient means of exploring and sometimes improving canal status.

### Electrodes

The use of four recording electrodes is required by the IHP.  Site preparation using alcohol pads and a mild abrasive gel is recommended.  Excessive abrasion must be avoided.  The non-inverting electrode is placed on the high forehead as close as possible to the hairline, in the midline.  Inverting electrodes are placed on each mastoid area, and this develops two differential recording channels: forehead to L mastoid and forehead to R mastoid.  A common or ground electrode is placed elsewhere on the head, typically on the forehead but not immediately adjacent to the non-inverting electrode.  The exact location of the common electrode is unimportant. The use of a neck (C7 spinal) position for the inverting electrode is sometimes recommended but is NOT appropriate.  While such an electrode may yield large ABRs, there is increased noise to offset that advantage, as well as loss of laterality waveform cues that are crucial for BC ABR.

Electrode impedances can have significant effects on EEG quality and therefore on successful testing.  Absolute values should be less than 10 kilohms and preferably less than 5 kilohms.  The larger these impedances, the larger the amount of pickup of external electromagnetic interference and of artifacts from movement of the electrode leads.  Even more important is the symmetry of the two electrodes that form each differential pair.  These should be as similar as is possible with reasonable effort.  A difference of not more than 1 kilohm is a desirable target, two kilohms is acceptable and three kilohms is borderline unacceptable.

The effect of large impedance differences is to degrade the common-mode rejection ratio (CMRR) of the preamplifier, that is, to reduce its ability to block EEG noise components that are common to both electrode sites in a differential pair.  The amount of reduction in CMRR increases almost proportionally with the impedance difference.  The effect can be to have difficulty achieving a satisfactorily low level of EEG noise for ABR recording, despite the fact that the child appears quiet.  Careful attention to electrode impedance asymmetry is required.

Given reasonable efforts to achieve satisfactorily low and symmetrical impedances, testing may proceed despite less than ideal conditions.  The audiologist should document the impedance values and be alert to the possible need for larger averages and more frequent replication of records, should EEG conditions require it.

## Recording Parameters

Specific recording parameters for the various components of ABR testing are given in Appendix B.  A few general points can be made.  First, the recording bandwidth must be appropriate relative to the frequency spectrum of the ABR waveforms of interest and the latencies of the key ABR features.  Thus, the optimal high-pass cut-off frequency for ABR threshold estimation is lower than that for otoneurologic ABR testing and CM measurement.  Second, the analysis window must be long enough to encompass the target response, so it must be greater for near-threshold recording.  Third, for threshold estimation the stimulus repetition rate should be as fast as possible, given a required analysis window length, because wave V is relatively unaffected by high repetition rates, which yield more averaging per unit test time.  The fourth point relates to EEG artifacts.  It is important to utilize the artifact rejection systems appropriately, and as a rough guide the amplifier gain should generally be set such that ‘quiet’ EEG occupies at least half of the vertical display range of the ongoing EEG.  This ensures that when EEG artifacts occur, the sweeps are likely to be rejected.  It is poor practice to record ABRs with the gain such that the ongoing EEG display is virtually flat.  In that condition, even a few isolated artifacts will be sufficient to degrade an average seriously, and this leads to unnecessary inefficiency of averaging and increases the likelihood of interpretive error, especially for threshold estimation.  An initial gain setting of at least 150,000 is generally appropriate.

The required values for the ABR test components are set up in parameter files that will be supplied to audiologists conducting IHP Assessments.  These files will be updated from time to time, as the need for enhancements to the test protocols is determined.  Because all IHP Assessments must be conducted under standard parameter conditions, local changes to test parameters are not permitted for Assessments funded by the IHP.  Audiologists are at liberty to use any parameters they see fit for non-IHP measurements, provided that those parameters are set up in non-IHP parameter files.  Caution is required because the number of protocol files distributed by the IHP may increase over time, and over-write local files.

## ABR Detection Judgements

Accurate and reliable judgement of response presence or absence is the key to ABR threshold estimation.  For a useful discussion of good practices and sources of error, see Stapells (2000).  At present, there are few commercially available statistical response detection algorithms for toneburst ABRs in a diagnostic assessment context, in contrast to the situation for click screening, for which several algorithms are commercially available.

Response detection judgements depend on appropriate averaging strategies, effective methods of interpreting averages, and the ability to decide whether the EEG conditions permit useful decision-making.  The minimum recommended unit of averaging for any given stimulus condition is two averages of about 2,000 sweeps each. In cases of uncertainty about response presence or absence, increasing the number of sweeps to 4000 and liberal use of a third average are recommended.  The routine use of more than three averages for a given stimulus condition is not recommended, because it is generally more effective to increase the stimulus level, where feasible, given persistent uncertainty about the response.  Similarly, the routine use of larger averages than about 4,000 sweeps is discouraged because it is inefficient, due to the law of diminishing returns within averaging.  To double the response to noise ratio in an average of 1,000 sweeps would take 4,000 sweeps.  To double it again would take 16,000 sweeps, which is impractical.

The interpretation of averages itself relies on four main cues: the occurrence of a response-like waveform at a reasonable latency, the prominence of that feature relative to the fluctuations in the remainder of the average, the reproducibility of the feature across averages, and the growth pattern of the feature over the course of individual averages.

The morphology of the ABR to tonepips is very different from that usually encountered in otoneurologic ABR testing with click stimuli.  The waveform changes are most marked near threshold and for stimuli at 0.5 kHz.  Typically, the earlier waves of the ABR are absent, and the response is a slow and late V-V’ negative-going transition.  There may be no wave V at all, but only a negative V’ peak.  There may also be a positive-going deflection following V’, at the end of the analysis interval.  The tonepip response at 2 kHz usually shows a wave V that is more clearly defined and sharper than at 0.5 kHz, and the 4 kHz response can be quite similar to conventional click responses.  At moderate sensation levels there may even be earlier ABR waves.  However, this will not usually be observed because the ideal IHP intensity tactics will direct effort to near-threshold recording conditions, to maximize test efficiency.

With regard to latency, in the Assessment context the range of possible latency is so broad that using specific targets for latency is unhelpful.  In general, the putative response should be larger than any other physiologic fluctuation in the average.  It should also demonstrate marked similarity across averages, and the level of similarity in the putative response region should be greater than elsewhere in the averages.  If a response is genuine, it will tend to develop steadily and progressively over the course of averaging.  For this reason, very close monitoring of the increasing clarity of the putative response is necessary.  Waveforms that develop suddenly at any point during the course of averaging should be regarded with suspicion.  In contrast, genuine responses can be rapidly obscured by small numbers of high-noise sweeps; this may happen frequently if the gain is insufficient relative to the artifact reject level, which is fixed at full scale on the ongoing EEG display of the current IHP equipment.

The ability to judge when EEG conditions are acceptable is extremely important in ABR threshold estimation – far more important than it is in otoneurologic ABR measurements.  In threshold work, there is a natural tendency to attempt to make response detection judgements and derive threshold estimates even under adverse EEG conditions.  This must be resisted.  For tonepip ABR threshold testing, EEG conditions MUST be satisfactory, or the results will be incorrect and a significant clinical interpretive error may follow.  It is far better to suspend judgement than to make such an error.  If the cause of the adverse EEG conditions cannot be rectified, then the ABR test should be aborted and rescheduled with appropriate attention to the causes of the poor EEG conditions.  It may be necessary to resort to testing under sedation, in these circumstances.

EEG quality judgements are presently subjective, and are a matter of training and experience.  Software that provides a quantitative statistical aid to the judgement of EEG quality is under evaluation and it is likely that numerical targets for EEG noise levels in averages will be available shortly.

## Definition of Threshold

There are several possible definitions of threshold arising from ABR measurements.  The current IHP definition is necessarily driven by the approach used in collection of the published normative data.  For any given frequency and route of stimulation, the value required to be reported is the ‘dial’ value on the ABR instrumentation (Biologic NavPro) at the lowest intensity for which a clear and reproducible ABR waveform is visually identified.  Where that value is the lower limit required by this protocol, it is the value to be reported.  Where the response level is greater than the lower limit, there must be a clear response-negative record at no more than 10 dB below the level reported as threshold.

The thresholds derived by ABR measurements have an indirect and statistical relationship to true hearing levels, and also to the target disorder.  ABR thresholds are conventionally expressed in dBnHL and are NOT generally equal to perceptual thresholds in dBHL.  An offset adjustment for bias of ABR thresholds is required.  The adjustment is typically between -5 and +15 dB, and may vary according to the type, frequency and severity of the hearing loss, as well as subject age and ABR testing procedures.  The evidence review to date indicates that provided the protocol defined here is followed closely, a constant adjustment that is specific to each test frequency and route of stimulation will yield acceptable threshold estimates.

The adjustments will be built into equipment calibration by the IHP in such a manner that ABR thresholds as defined in this document will equal the numerical values in dB of the required perceptual threshold estimates.   These values will be referred to as dB ‘estimated hearing level’ or ‘dB EHL’, as a reminder that they are not defined in conventional dBHL or dBnHL.  Given this approach, the target disorder is equivalent to ABR thresholds of greater than 25 dBEHL, and the ABR thresholds in dBEHL may be used directly in any subsequent prescription for amplification.

# Sedation

There is very extensive experience from Ontario and elsewhere indicating that in an infant under about six months of age, tonepip ABR testing can almost always be done satisfactorily with the baby in natural sleep.  Appropriate training and test protocols are necessary and sufficient.  Routine resort to sedation (or general anesthesia) in infants under six months corrected age is not recommended.

Testing under sedation may be necessary in infants for whom acceptable behaviour and EEG conditions cannot otherwise be obtained.  Usually, at least one attempt to test in natural sleep would have failed before resorting to sedation.  It is reasonable to consider fairly routine use of sedation in children older than six months or for children who have to travel long distances for Assessment, such that it is especially important to have a reasonable assurance of success.

The audiologist determines that sedation is indicated on audiometric grounds. The family determines whether sedation will actually occur, in consultation with the audiologist and appropriate physicians.  The infant’s pediatrician or family physician would normally be involved, as he or she may have unique knowledge of contraindications or risk indicators in the history.   Where specific centers have established, high-quality protocols in place, they should take precedence.  Documented, informed consent would normally be required.

If sedation is indicated and consented, a physician should prescribe an appropriate agent (usually chloral hydrate) and be accountable should any complications arise. Appropriate risk management procedures to guard against adverse events such as respiratory depression should be in place.  While there is wide variation in practices for sedation with chloral hydrate, the IHP recommends adherence to a conservative standard of care.  Testing under sedation should normally be done under medical order and preferably with medical or nursing supervision of the infant from the time of administration through to the end of the indicated recovery period.  Immediate access to emergency services may be sufficient, depending on local safety protocols.

In a few infants, especially those with neurological and/or behavioural disorders, the response to sedation may be paradoxical activation.  This has been addressed in various ways, by increased sedative dosage, by use of alternative medications, or by resort to light, general anesthesia.  The indications for these procedures are a matter for local risk management protocols and standards of care.

While testing under sedation is generally easier than under natural sleep, not all IHP audiologists will have access to the required medical coverage. The decision whether or not to accommodate testing under sedation rests with the individual audiologist.  Where necessary, cross-referral may occur to another IHP audiologist who has a sedation practice. The IHP will attempt to facilitate reasonable access to testing under sedation, at least at a regional level.

# Calibration

The factory calibrations supplied with the IHP instrumentation are not acceptable for Assessments and shall be modified according to IHP specifications.  Many factors affect the proper calibration of stimuli for tonepip ABR testing in infants.  Important variables include subject age, stimulus route and transducer, frequency, envelope and repetition rate, the amount of averaging, EEG criteria for residual averaged noise, and definition of threshold.  The amount of high-quality published data is limited but sufficient to establish normative calibrations.  Provided that the key features of IHP test protocols match those of the protocols used in deriving the published data, the use of published norms is far superior to the development of local norms.

The IHP will be responsible for the determination and dissemination of reference settings for the calibration files.  Normally, calibration files or values will be distributed by e-mail. In February 2002, initial calibration values in Estimated Hearing Level (EHL) intended for routine clinical use were distributed, with a description of their derivation and use.  These values will be updated from time to time, as further information becomes available.

Because of the digital nature of modern ABR equipment, the main reason for routine, full acoustical calibration is mechanical change in the stimulus transducer.  It is planned to conduct annual calibrations, to ensure correct transducer function.  Calibration services will be arranged by the IHP, with reasonable notice to the Assessment centres. The individual IHP audiologist is responsible for routine, daily listening checks, and these are a procedure required by the IHP.

The maximum stimulus levels to be used in the Assessment are a matter of judgement by the individual audiologist. The IHP is unaware of any adequate scientific evidence on which to base derivation of valid damage risk criteria for the infant cochlea.  The maximum SPL values developed are typically about 120 dB, and such values have been used in ABR test protocols for many years.  In ABR testing, the stimuli are intermittent and there is no exposure to high levels of continuous stimulation.

## Clinical ABR Records

General requirements for record-keeping were noted earlier.  The required deliverables from testing under IHP funding include the brief, summary report to the PSL co-ordinating agency, as well as adequate clinical records held locally.  For ABR testing, a complete, hard-copy record of the ABR test results, including waveforms, shall be retained by the IHP audiologist in the infant’s clinical chart.  The ABR source data files shall also be retained electronically. All IHP records for any infant must be readily accessible, given the name and date of birth of the infant.  For any infant, the records for any specified test date must be similarly accessible. The records for any given test must permit identification of the stimulus and recording parameters for each individual average waveform, as well as the order of acquisition of averages.  Specifically, the stimulus type, level and route, the number of sweeps averaged and the laterality of the average must be identifiable.  For BC records, two-channel (bilateral) recordings must be available and their laterality identifiable.

# Threshold ABR with Tonepips

At present, it is widely acknowledged that tonepip ABR is the method of choice for frequency-specific threshold estimation in young infants.  While the click ABR is a convenient tool, there are many reports of substantial audiometric errors using click stimuli.  Consequently, being evidence-based, the IHP cannot consider click ABR threshold testing alone to meet a generally acceptable standard of audiometric assessment quality.  Similarly, for newer tools such as single-frequency or multi-frequency, steady-state evoked potentials (SSEPs), published evidence about audiometric accuracy is currently considered insufficient to justify clinical implementation within the IHP.

Tonepip ABR is specifically recommended by the US Joint Committee on Infant Hearing (2000).  There is reasonable evidence from meta-analysis that given appropriate test protocols, tonepip ABR thresholds by air conduction can predict conventional audiometric thresholds with reasonable accuracy (typically within about 10 dB) for a wide variety of hearing losses in adults and children (Stapells 2000).  Therefore, for IHP Assessment, tonepip ABR is a required procedure for puretone threshold estimation in infants under six months of age.  It is also recommended strongly for use in older infants for whom VRA is deemed unfeasible or unreliable, and where VRA does not provide sufficiently ear-specific and frequency-specific information to satisfy the audiometric objectives of the IHP.  Testing under sedation may be necessary, especially in older infants.

The technical parameters for successful tonepip ABR recording are given in Appendix B.  Parameter files will be defined and distributed by the IHP.  The parameters are tightly coupled to calibration values and norms derived by critical review and meta-analysis of published data.   Because the parameter selections are strictly based on high-quality evidence, they may differ from various popular recommendations.  Examples of this include the use of alternating tonepip polarity, and the use of tonepip envelope plateaux.  Such differences have been carefully scrutinized and found to be justified.

While the frequency-specificity of tonepip ABR threshold estimates may be improved slightly by an ipsilateral masking method such as notch-masking, the likelihood of clinically significant errors in threshold estimation if such masking is not used is low.  Also, it is possible that response depression due to physiologic masking spread may occur at high stimulus levels.  While the judicious use of such masking may improve results in a few cases, there are additional concerns regarding technical aspects, test complexity and consistency across Ontario.  Therefore, the use of ipsilateral masking noise is not recommended within the IHP protocol.

The frequencies at which tonepip ABR testing shall be attempted include at a minimum 0.5 and 2 kHz.  Both are important generally and also particularly with regard to prescription of amplification.  If there is severe or profound impairment at 2 kHz, 0.5 kHz is a key measure of residual low-frequency hearing.  If reliable thresholds have been obtained at 0.5 and 2 kHz, the clinical utility of measurements at 1 kHz depends on the 0.5/2kHz threshold difference.  If the difference is less than 30 dB, interpolation of the 1 kHz threshold as the average value will rarely be seriously in error.  However, if the difference is 25 dB or greater, then threshold measurement at 1 kHz is recommended.  The decision in the individual infant is at the discretion of the IHP audiologist.

ABR measurements at 4kHz are recommended if the ABR threshold at 2 kHz is more than 10 dB above the minimum necessary test level (typically about 25-30 dBEHL, or 30-40 dBnHL).  They are recommended also if the 2 kHz threshold is within normal limits but DPOAE results suggest an abnormality above 2 kHz.   The 4 kHz threshold carries some limited information for immediate management, but it may be a risk indicator for progressive impairment and the need for ongoing monitoring.  The decision to test by ABR at 4 kHz in the individual infant is at the discretion of the IHP audiologist, weighing the clinical need and feasibility of obtaining useful results.

## Overall Stimulus Strategy

Many infants presenting for Assessment will have normal or near-normal hearing.  The initial key question is whether the target disorder is present.  The minimum required level for testing is currently defined as 30 dBEHL at all frequencies.  In the absence of prior audiometric data, testing should start at 2 kHz because it is clinically the most important single frequency. The choice of initial test ear should match the AABR screening result, if available.  If referral were bilateral, the ear of convenience would be chosen.  The AC route is used first, because it is overall hearing sensitivity that is important initially.

If the starting ear gives a clear and reproducible 2 kHz response at 30 dBEHL AC, the key question is now whether the other ear is normal at 2 kHz.  Therefore, where practicable the other ear would now be tested at 2 kHz, 30 dBEHL AC.  A normal result bilaterally would be followed by exploration of 500 Hz.  A questionable response at 30 dB might justify an intensity increase to 40 dB, but no response at 30 dB 2kHz AC could be followed either by a BC test at 30 dB or perhaps more effectively by a continuation of AC testing at, say, 60 dB.  No response would establish a substantial hearing loss, and the key question is now about type of loss, so a 30 dB BC test is indicated.

There are many combinations of possible response outcomes, and many factors that will influence the precise choice of next stimulus condition in the individual case.  These are at the discretion of the IHP audiologist, who is in a position to weigh all the factors.  Usually, efficient strategies targeting key clinical information tend to involve more switching of ears and changing routes of stimulation than may have been the testing pattern traditionally, for older infants.  One issue is whether this approach is physically practicable and is likely to awaken the child.  Clinical experience and judgement are required here, but it is stressed that conventional tactics may be inefficient and practices that may be necessary in older infants may be inappropriate in infants under three months.  Infants under about three months sleep more readily and with less dependence on position than older infants.  At two months of age, the typical infant preferably will be tested laying supine.  In that situation, changing ears and transducers is easier than if testing were done in the caregiver’s arms.

Efficient testing strategies take account of probable patterns of hearing impairment and implications of results for further action.  A common example is the situation in which there is an early indication of a conductive component from MEA or from an air-bone gap, or both.  In the absence of a known, structural abnormality, the most probable cause is otitis media.  In that case, precise estimation of AC hearing thresholds over several frequencies may be unnecessary, because the hearing is likely to change over time or following medical intervention.  An emphasis on reliable demonstration of a substantial conductive component and determination of BC thresholds seems a reasonable and sufficient approach.  Precise quantification of the air-bone gap in this situation may be an inefficient use of valuable program resources that could be directed more effectively, perhaps to re-assessment of infants with proven PCHI.

# Stimulus Level Tactics

The IHP protocol is based on direct estimation of thresholds and specifically does NOT invoke estimation based upon latency input-output functions, for which there is evidence of insufficient accuracy and reliability.  Therefore, the general use of intensity series with small step size is specifically discouraged because it is relatively uninformative and inefficient.

The detailed tactics of selecting stimulus intensity levels have a large effect on overall test accuracy and efficiency.  The most common cause of inefficiency is to use a step size that is too small, or to fail to make use of prior information in setting levels that will be close to threshold.  The optimal result for any particular frequency and route of stimulation is to bracket threshold with only two levels that are no more than 10 dB apart, in which case only two pairs of averages are needed to define the threshold.  Skilled testers frequently can determine the threshold with only three pairs of averages, that is, by measurements at only three intensity levels.  The more intensity levels required, the less efficient the strategy.  An approach for selecting intensity levels that is generally efficient for the early stages of threshold estimation in the absence of prior knowledge is to approximately bisect the current range of intensities in which the threshold is believed to lie.  Such a strategy yields important clinical information very quickly.  For example, if there is clearly no response at 30 dB, say, then an efficient next level is 60 dB.  This approximately bisects the current range of uncertainty about the location of the true threshold (30-90+), and is unlikely to disturb the infant.  The outcome at 60 dB is much more informative clinically than an outcome at, say, 40 dB would have been.

In general, the most efficient intensity tactics start with large step sizes, converge quickly and finish with the minimum necessary step size to resolve threshold to acceptable precision.  A 10 dB maximum separation of the response-positive and response-negative levels is acceptable.

## Multichannel Recording

The four recording electrodes are configured as two EEG amplification channels, ipsilateral and contralateral to the stimulated ear.   For AC stimuli, the contralateral channel has limited value and need not be retained in the audiologic records.  It is sometimes helpful on-line in identifying errors of stimulated ear selection.

When stimulating by bone conduction, the two channels are necessary in order to resolve the responding cochlea; wave V is earlier and generally larger in the channel ipsilateral to the responding cochlea.  This approach avoids several practical difficulties and unsolved questions related to contralateral masking of BC ABRs in infants.  The use of two recording channels in BC testing is a required procedure.

Click ABR Measurement

In the IHP context of audiometry based primarily on tonepip ABR by air and bone conduction, ABR measurements using clicks have a secondary role, but an important one.  There are three aspects of diagnostic inference that may be clarified by the use of click stimuli.

## Retrocochlear Pathology

The clinical implications of indications of retrocochlear abnormality in the infant ABR are not well-established in early infancy, and a detailed exposition of otoneurologic outcomes and interpretation is beyond the scope of this protocol.  However, it is the accuracy of tonepip ABR threshold estimates that is the crucial issue here, and inaccurate threshold estimates may arise because of neuronal dys-synchrony that causes reduction or absence of the V-V’ complex.  This may be due to a disorder affecting the integrity of neuronal response from afferent auditory brainstem pathways.  Such disorders are not common in young children but include space-occupying lesions (such as in neurofibromatosis), neurodegenerative disorders, structurally-mediated disorders (such as in hydrocephalus) and vascular disorders.  In the presence of such conditions, using click stimuli at a high intensity it is often possible to observe an ABR morphology that suggests a retrocochlear location of disorder.  Usually this involves presence of one or more early waves, prolonged interwave intervals and/or depression or absence of later ABR waves.

When a clear and reproducible wave V or V’ is seen at any stimulus level in the course of tonepip ABR measurements, significant audiometric inaccuracy due to a retrocochlear disorder is very unlikely.  Therefore, it is not considered necessary to measure click ABRs routinely.  However, if there is no clear wave V-V’ at the highest tonepip levels available, an ABR measurement using a click at or near the highest level available (typically about 95 dBnHL) is required.  Modifications to the stimulation and recording parameters for this situation will be distributed as an addendum.  Separate measurements of response to condensation and rarefaction clicks are required.  If either record individually shows a clear wave V, the likelihood of substantial inaccuracy in the tonepip thresholds caused by retrocochlear disorder is low.  If a depressed or absent wave V is seen in BOTH click records, then confidence in the accuracy of tonepip threshold estimates is reduced.  If no ABR wave I through V are seen, then the interpretation depends on whether or not there are indications of possible auditory dys-synchrony, as outlined shortly.  A clear finding that suggests a retrocochlear disorder should be noted in the report, and clinical actions following that finding, such as a medical referral, are at the discretion of the audiologist.

If early waves but no wave V are seen, one possibility is that there is a retrocochlear disorder.  The other is that the early waves are not neurogenic but are cochlear microphonic potentials (CM) or stimulus artifact.  These alternatives can usually be discriminated by the use of separate averages with condensation and rarefaction clicks.  If the early waves do not invert, they have a neuronal origin.  If they invert, they are likely to be either CM or stimulus artifact.  To rule out stimulus artifact, the stimulus delivery tube is disconnected from the ER3A transducer (NOT from the ear insert) and its end moved away from the transducer without changing the position of the ER3A.  If the early waves remain, then they are due to stimulus artifact and there is no reason to question the reliability of the tonepip ABR thresholds.  If the early waves are CM, then confidence in the tonepip thresholds is decreased, just as it would be if the early waves were of neural origin.  The finding of a clear CM and absent neurogenic waves suggests auditory dys-synchrony, and the validity of ABR thresholds as estimators of perceptual thresholds is highly questionable in that case.

## Auditory Dys-synchrony

Auditory dys-synchrony (AD) is defined by a cluster of findings that includes normal or near-normal OAEs, absent or severely abnormal ABRs, absent middle-ear muscle reflexes, and actual hearing impairment of any degree from mild to profound.  The disorder is almost always bilateral.  In a small proportion of cases the OAEs are absent or degrade over time.  In older children there may be other concurrent peripheral neuropathies.  AD may represent as much as 5-10% of all sensorineural hearing impairment in infancy, but its actual prevalence is not yet well-understood.  There are several underlying variations, mechanisms and causes, at least some of which appear to be genetic.  Hyperbilirubinemia is also a risk indicator.  In the presence of AD, ABR threshold estimates in infants are not reliable and currently it is not uncommon to delay intervention until at least some behavioural threshold information can be obtained.

It is a feature of AD that the CM are usually large and easily recorded, but there are usually no ABR waves in response to either tonepips or clicks.  Therefore, care must be taken to differentiate CM, stimulus artifact and neurogenic response, using the method noted above. The CM is readily apparent as a series of large deflections in the 0-5 ms latency region that should invert with change in click polarity.  It is currently believed that given an absent or highly abnormal click ABR, the presence of a large CM suggests a differential diagnosis of AD, especially when OAEs are absent.

The appropriate course of action, given a probable AD, is at the discretion of the individual IHP audiologist.  An ancillary protocol that will address additional diagnostic procedures, as well as management issues, is in preparation.  It is likely that cortical evoked potentials will contribute to any such protocol, which may extend to include those cases of retrocochlear disorder, wherein the tonepip ABR threshold estimates are also suspect.

## Other Causes of Tonepip/Click ABR Discrepancy

If a clear wave V is seen in response to clicks, in the absence of tonepip responses at maximum AC levels, the most common cause is that there is better hearing at some frequency not yet tested by tonepips, in the range 0.5-8 kHz that is excited by clicks.  This may be explored by tonepip ABR measurements at the untested primary frequencies, at least in the range 1-4 kHz.  It should be noted that stimulus calibrations are less dependable at 6 and 8 kHz.

An occasional cause of a wave V-V’ to clicks but not to tonepips is that there is severe cochlear impairment and insufficient synchronous excitation of primary neurons with a frequency-specific stimulus, whereas the click excites a broader region of the cochlear partition and with greater synchrony. In that situation, no tonepip response will be seen at any frequency, but the click threshold could be as low as 85 dBnHL (note that the EHL conversion is not applied for clicks, because click thresholds are not reliable estimates of perceptual thresholds).  If this pattern of results is seen, the audiometric values reported should be based on the tonepip results, but the finding of a click response should be noted and the possibility of slight overestimation of thresholds be taken into account in follow-up measurements.

Finally, if there is a marked, unexplained discrepancy between click and tonepip data, measurement error is the probable cause and all the results should be critically reviewed.

Measurement of RECD in infants identified with PCHI

Acoustic Characteristics of the Ear Canal

The acoustics of an infant’s external ears differ significantly from those of the average adult (Kruger, 1987). It is recommended that measurement of the Real Ear to Coupler Difference (RECD) is completed at the time of the diagnostic ABR in those infants with identified hearing loss. In the event that the individual RECD measurement is unobtainable, age-related predicted values can be applied (normative values are provided in Appendix E). When audiologists describe hearing loss in dBHL and apply these average adult values dB adult EHL to infants, errors in estimated hearing sensitivity can occur. This can be detrimental to the prescription of hearing aids for use by infants. It is recommended that if a child is diagnosed with PCHI and may be a candidate for amplification pending the parents’ decision, the RECD should be measured at the time of the diagnosis.

The information may also be useful in the event that over time, a shift in hearing sensitivity appears between ABR and VRA, or subsequent VRA assessments. Descriptions of procedures for adjusting assessment data to remove effects of ear canal resonance over time, and for combining data from multiple test sessions is available separately.

Procedure for Measuring and RECD

RECD Measurements should be obtained from each infant using SpeechMap® DSL® in the Audioscan RM500 hearing aid test system (Etymotic Design, Inc) following the procedure described by Moodie et al (1994). Briefly, the HA-2 coupler is connected to the coupler microphone of the unit and as ER-3A insert earphone transducer is coupled to the other end of the HA-2 coupler. A swept-frequency stimulus generated by the probe microphone system is delivered into the coupler and the coupler response is measured by the microphone. Couple an insert foam tip and RECD transducer and insert the foam tip into the infant’s ear. It may be helpful to couple the probe tube to the insert phone or probe tip with hearing aid moisture guard wrap (for very small ear canals). This is helpful in coordinating insertion and ensuring 2mm remain at the probe tip edge. The same stimulus is presented via the probe microphone system and insert phone coupling, and the response is measured. The difference between the real-ear response and the coupler response is measured. This difference is the individual transfer function designated as the RECD.

Specific instructions to obtain accurate RECD measurements with the Audioscan® RM500® and Verifit systems are in Appendix F.

Visual Reinforcement Audiometry (VRA) – General Considerations

Visual Reinforcement Audiometry (VRA) is an operant conditioning technique used with infants and young children in order to evaluate their hearing acuity. The child is trained to turn his/her head to look at an attractive visual object, such as an animated toy, in response to an auditory signal. The lowest intensity level at which the auditory signal reliably elicits this response is considered to be the child's Minimum Response Level (MRL).

VRA is part of the on-going Audiologic Assessment of infants registered with the IHP.  VRA can provide frequency- and ear-specific information about the child's behavioural responsiveness to sound.  The VRA measurements can help to confirm previous findings about the child's hearing and determine the stability of a hearing loss established through other assessment techniques.  It also can provide valuable audiologic information upon which hearing aid fitting and/or adjustment may be based.

Target Population and Referral Sources

Within the IHP, an infant previously identified as having a permanent childhood hearing impairment (PCHI) on the basis of the ABR-based assessment protocol would require behavioural audiometric assessment prior to reaching 12 months of age.  Infants of at least 6 months corrected age and up to approximately 24-30 months of age (or until the child can reliably perform play audiometry) are candidates for VRA. For premature infants, Moore, Thompson, Folsom (1992) suggested that "a corrected age of 8 months and/or a mental age of 6 months is typically required for acceptable performance using VRA".  Infants with developmental delay may not be ready for behavioural testing using VRA at 8 to 12 months of age, so information about an infant's cognitive developmental status may be useful in deciding when to schedule the infant for this test  (see section on Pre-test Preparation). VRA testing at a corrected age of 9 to 10 months for most infants is considered to be optimal.

Infants may be referred for VRA testing through the following routes or in the following circumstances:

- From IHP screening and assessment, where ABR results indicated a permanent hearing impairment of >25dB EHL at any frequency in the range of 0.5-4kHz, in either ear.

- From IHP screening and assessment, where the infant has reached the corrected age of 8 months, but ABR results have been incomplete, inconclusive or unreliable, and a hearing loss has not yet been confirmed or ruled out.

- From IHP re-screening in the second year of life (age 12 to 24 months) due to perinatal indicators and where results suggested a possible hearing impairment.

- As part of the hearing aid evaluation and follow-up process dictated by the IHP Amplification Protocol for infants identified with PCHI who are using assistive technology.

- From physician referral of children up to 24 months of age who have acquired or who demonstrate any of the risk indicators for PCHI (Appendix A)

- From physician referral of children up to 24 months of age who have acquired or who demonstrate meningitis or significant head trauma as risk indicators of PCHI.

Test Personnel

It is required that two testers be involved in the VRA test - one as examiner who presents the stimuli and visual reinforcement and records the infant's responses, and the other who is in the room with the infant and acts as distracter.  The examiner must be an audiologist who is registered with CASLPO and who has attended an IHP VRA training workshop.

The role of distracter may be assumed by another audiologist or by an individual who is trained and supervised by the examining audiologist.  In some instances, the parent may be used in this capacity, at the discretion of the audiologist.

Test Environment

Testing shall be done in an audiometric test room satisfying standard criteria (ANSI S3.1-1999) for maximum permissible ambient noise for audiometric test rooms.  The room shall be of sufficient size to accommodate the parent, infant and distracter comfortably and shall be adequately ventilated. The room shall contain minimal visual distractions to the infant.  It is recommended that the room lights be provided with a dimmer switch, to allow for enhancement of the illuminated visual reinforcers.

The presentation of stimuli and visual reinforcement are controlled from a second (observation) room.  The test and observation rooms are separated by a one-way window so that the infant is not distracted by the examiner, but the examiner can see the infant clearly.  (Darkening the observation room may help to achieve this condition.)  Two-way communication must be available between the examiner and the distracter.

Test Equipment

VRA testing requires a clinical diagnostic audiometer that meets the ANSI S3.6-1996 specifications.  The audiometer shall be capable of presenting pure tone and FM warbled-tone stimuli through insert earphones, supra-aural earphones, and a bone conduction oscillator. Loudspeakers could also be available for those children who will not tolerate earphones initially, for hearing aid validation using a behavioural response, and for counseling and educational purposes. Warbled tones and are the stimuli of choice for the IHP protocol.

Stimulus Transducers

In the absence of specific contraindications, insert earphones (ER-3A) are the required transducers for air-conduction VRA testing.  Insert earphones have several advantages over supra-aural earphones for this test method, including reduced acoustic crossover (increased inter-aural attenuation), decreased likelihood of collapsed external ear canals, accurate location of sound delivery, increased comfort and reduced interference with head-turn response.  Research has shown that tolerance of insert earphones by infants is generally good once the earphones have been placed in the ears (Widen, 2000).

Supra-aural earphones (TDH 49/50) are to be used when insert phones are contra-indicated, such as when the ear canals are very small or stenotic or when the infant does not tolerate the insert phones.  Careful attention to accurate placement of a TDH earphone is especially important to ensure appropriate stimulus levels and to avoid collapsing ear canals.  Soft padding between the headband and the top of the infant’s head may by needed to ensure comfort and proper placement.

A bone vibrator with a plane circular tip “with a nominal area of 175 ( 25 mm2”, as specified by the ANSI S3.6-1996 Specification for Audiometers, is required.  Bone conduction thresholds are needed to identify and quantify conductive and sensorineural components of a hearing impairment.  Establishment of these thresholds requires accurate and stable placement of the bone oscillator.  Soft padding under the headband may be helpful.  If proper force and stability of the bone conductor cannot be achieved with the standard headband, a band of elastic fabric with Velcro attachments may be used (see ABR section on Stimulus Transducers).

Visual Reinforcement

A toy located in a smoked Plexiglas box that can be illuminated and animated is the recommended visual reinforcement.  At least two such toys, one on each side of the room, are required; four toys - two on each side - are preferable.  The toy in the box should not be clearly visible to the child unless the box is illuminated.   A switch in the observation room controls the animation and bright illumination of each individual toy.   The reinforcement toys should be non-aversive to the child. It is recommended that the animation and the illumination of the toys be independently controllable, as some children are afraid of the animation, and thus are more relaxed if only the illumination is provided when the reinforcement is presented.

Toys for Distracting

A large supply of different toys is needed to distract the child.  These should be toys that can be manipulated quietly, can be effectively cleaned and/or disinfected as necessary, and are not similar to the reinforcement toys. These distracting toys should be less interesting and appealing to the child than the reinforcement toys.

Calibration

The ambient noise in the test room must meet the ANSI S3.1-1999 criteria.

Calibration of insert earphones, supra-aural earphones and bone vibrator must be carried out according to ANSI standards (ANSI S3.6-1996).   It is required that this calibration be done when the audiometer is installed (or moved), and then annually thereafter.

A visual examination of the equipment and a listening check, at all frequencies being tested, shall be carried out at least once a month.  It is also necessary to do a brief listening check of at least one frequency in each transducer before each test.

Test Set-up

In the test room, the infant is seated on the parent's lap, gently supported and facing forward.  If appropriate, the infant may instead be seated in a highchair (that meets all pertinent Canadian safety standards), or an older child may be seated on a low chair, with the parent seated beside and slightly behind the child.  A low table may be placed in front of the child to provide a surface for the distracting toy or activity.

The distracter is seated on a low chair on the other side of the table and facing the child, with a concealed collection of toys available to be used to maintain the child's attention between stimulus presentations.

The reinforcing toys are at 90 degrees to the infant's right and left side.  The reinforcing toys should be at the child's eye-level, to maximize their visibility and simplify the head-turn response. (See diagram, Appendix D) The angle at which the reinforcers are placed may be reduced somewhat if the child is not physically able to give a full head-turn response.

The audiologist in the observation room must be able to see clearly the child's face and the distracter's activity.  The audiometer, reinforcement controls and recording materials must be easily accessible. There must be good two-way communication between examiner and distracter.  However, the examiner's voice should not be audible to the child unless it is being presented as a speech stimulus.  The distracter in the test room may use an earphone, headphones or a bone conduction oscillator wired to the audiometer, or an FM system may be used for examiner-distracter communication.

Pre-test Preparation

Several issues should be discussed with the parent or guardian at the time that the VRA appointment is being made.

- Because the VRA technique requires that the infant be attentive to sensory stimulation during the test, the VRA appointment should be scheduled for a time of day when the infant is most likely to be alert.

- The parents should be asked about the medical condition of the baby's ears.  If possible, the infant should be tested when s/he is clear of ear infection.

iii) The parents should be asked about the status of the infant's vision.  If the infant has visual problems, the test environment may have to be adjusted (reinforcers moved closer to the infant, room lights dimmed).

iv) Information about the infant's developmental status can help to determine whether the VRA method is likely to be appropriate for the particular infant.   For example, if the baby, at 8 months, is unable to sit up or turn his/her head, VRA should probably not be attempted until s/he has better motor control. However, many infants with some degree of developmental delay can perform quite reliably in VRA.

The test procedure should be explained to the parents when they arrive for the test.

Otoscopy

It is recommended that, at the beginning of the test session, the audiologist conduct a cursory otoscopic examination.  If the ear canal is obviously occluded, removal of the occlusion by an experienced physician is recommended.  However, successful testing can often be accomplished in the presence of substantial debris or cerumen. The decision to undertake testing with insert earphones under such conditions is at the discretion of the audiologist.  If the insert becomes blocked, it may be found that probe removal withdraws the debris/cerumen and that reinsertion leads to a successful test.

If insert phones cannot be used, supra-aural earphones are an option, with the caveats noted earlier concerning placement.  It is also possible to undertake bone-conduction testing, but a return visit for AC testing would usually be required to demonstrate normal hearing.

Partial occlusion is not uncommon, but usually does not interfere with accurate testing. VRA should be attempted in all but the most severe cases of apparent occlusion or draining ears.

VRA Test Protocol

The objective of the ABR Assessment is to establish thresholds at a minimum of two frequencies bilaterally, in order to provide an assessment that is both expedient and accurate. In cases where VRA assessment is expected to confirm results obtained via ABR, the aim is to establish minimum response levels (MRLs) for air-conducted tones in each ear for at least two frequencies (namely, 500 and 2000 Hz), and to establish a bone-conduction MRL for at least one of these two frequencies at which there is an air-conduction MRL of >30 dB HL in both ears.

In most cases, the VRA Assessment will provide more comprehensive information about the infant’s behavioural responses to sound and it should be possible to obtain MRLs for at least three frequencies (in the range of 500 to 4000 Hz) bilaterally.  Consideration should be given to the following:

- The clinical relevance of the MRL at 1000 Hz increases with the increased difference between the MRLs at 500 and 2000 Hz.

- 3000 or 4000 Hz may be more important for hearing aid fitting than 1000 Hz.

For the purpose of this test protocol, MRLs of 30 dB HL or less are considered to be within normal limits.  (See section on interpretation of test results.)

The audiologist may determine that, in some cases, presenting the stimuli in sound-field is helpful in conditioning a child who is initially reluctant to wear insert earphones.  In this context, sound-field stimulation has a role to play in the VRA process.  However, as sound-field testing does not provide ear-specific information, sound-field results are not considered a sufficient basis for quantitative audiometry or a basis for the prescription or verification of assistive technology. Ear-specific audiologic information is required to meet the objectives of the VRA Assessment.

# Order of test stimuli

For air-conduction testing, pulsed, FM warbled tones of 1-2 seconds duration are presented through an insert earphone.

If the examiner has previous hearing threshold information about the infant, it is recommended that the testing begin with the better ear, and with the frequency at which the hearing is the most sensitive, based on the previous assessments.  This helps to ensure that the infant is conditioned to the VRA task as efficiently as possible.  If no previous information is available, or if previous information suggests no significant difference among frequencies or between ears, it is recommended that the air-conduction frequencies be tested in the following order:

- 2000 Hz, followed by 500 Hz, in the first ear

- 2000 Hz, followed by 500 Hz, in the second ear

- 3000 or 4000 Hz, at the discretion of the audiologist and based on diagnostic and/or amplification requirements, in each ear

- 1000 Hz, in each ear

2000 Hz and 500 Hz are tested first, in order to judge the stability of the hearing loss based on the ABR results, and because of their importance in speech perception.

If there is a significant ((30 dB) difference between the MRLs for 500 and 2000 Hz, 1000 Hz could be tested before 3000 or 4000 Hz.

A bone-conduction MRL shall be established for at least one frequency where there is a bilateral MRL of  >30 dB HL.  The bone conductor shall be placed on the side that has the lower (better) air-conduction MRL at the test frequency.

An air-conduction MRL for Speech Awareness may be established for each ear at the discretion of the audiologist, provided that this does not compromise the achievement of the IHP goal of establishing MRLs for frequency-specific stimuli.  Speech stimuli may be used during initial conditioning if the infant does not respond readily to tones, or may serve to regain the infant’s attention if it is waning after several warbled-tone frequencies have been tested.

# VRA Test Procedure

The recommended VRA protocol for determining Minimum Response Levels is based on the procedure described by Widen et al. (2000).

Once the infant and parent are seated in the test room, the audiologist shall explain the test procedure to the parent, and emphasize to the parent the importance of not cueing the child to the signals, and of not causing distracting noise.  The foam tips attached to the insert earphones are placed in the infant’s ears, with the other end of the tubing clipped to the back of the infant's clothing.  If the baby is irritated by the earphones and tries to take them out, it might be advisable to have an earphone in only one ear at a time during testing.  One method of keeping the infant from pulling the earphones out is to have him/her hold a small toy in each hand.   Supra-aural earphones may be used if the infant's ear canal(s) is extremely narrow or atretic.

Pulsed, warbled tones of 1-2 seconds’ duration shall be used as the initial stimuli.  The inter-stimulus interval (ISI) should be varied, and initially lengthened if random head-turns are frequent. (Allow about 4 seconds for the response, if it occurs, followed by reinforcement and then a 4-second interval before the next trial.)  A 10-dB step size is used (20 dB down, 10 dB up), in order to reach the MRL quickly.   A short, live-voice speech signal, such as "Hello there!", "Ba-ba-ba" or "Hi , …infant's name" may be used during conditioning, if the infant does not respond well to the tones.  Presenting the stimuli in sound field  through loudspeakers may also be used during conditioning if the infant is initially reluctant to wear earphones, however it should not be used as part of the Assessment as it will not provide ear-specific information.

## Role of the examiner (audiologist)

The audiologist, seated in the observation room outside of the test room, controls the presentation of stimuli and reinforcement.  S/he observes the infant’s behaviour, judges if the infant has responded or not, and records these judgments on the VRA Worksheet.  S/he may also guide the distracter in keeping the infant in the appropriate state for testing.

Role of the distracter

In order for VRA to be successful, the infant must be attentive to the task of listening for the sound that signals that there is a visual reward to follow.   The role of the distracter, who is in the room with the infant, is to keep the infant in an appropriate state of listening for the sound stimulus, while keeping the baby’s gaze toward the midline, so that the full 90-degree head-turn response can be seen easily when it occurs.

The distracter must maintain a fine balance between being sufficiently interesting to hold the infant’s visual attention, but not so entertaining that the infant ignores the sound signal and/or the visual reinforcer. The distracter must avoid making any noise (e.g. by talking or through noisy manipulation of the distracting toys) which could be mistaken for a stimulus by the infant or which could interfere with the infant’s ability to hear the sound stimulus.  However, it may be appropriate for the distracter, from time to time, to give the infant some social reinforcement, such as saying “Good work!” or “Good girl/boy” when the infant has responded correctly.  The amount and type of distraction activity needed to maintain an infant in a listening state will vary from infant to infant.  For one infant, the mere presence of the distracter may be enough to keep him/her facing forward, another may need to have his/her attention brought back to midline after each response.  The reliability of the responses may be a good indicator of whether or not the infant’s attention to the task is being maintained.

During the testing, the distracter must avoid giving the infant any cues that a sound signal has occurred or that a control trial is being carried out. Therefore, the distracter must be careful not to change his/her ongoing activity when a trial begins.

Another role of the distracter is to establish and maintain the cooperation of the parent during the test session.  The distracter helps the parent to maintain the infant in a good listening position; usually the baby responds best when seated, facing forward, on the middle of the parent’s lap in an upright position, not leaning back.  The distracter instructs the parent to keep the baby’s hands away from his/her ears, so that the earphones are not dislodged or removed.  The distracter reminds the parent to be as quiet and as non-interactive with the infant as possible during the test so that the infant attends to the task.  It is important that the parent avoid cueing the baby to the presence of the sound signal.  Sometimes a parent will respond with subtle body language, such as holding the baby a little tighter, or moving slightly, when the sound is presented.  The parent may do this subconsciously, and it is the job of the distracter to point this out to the parent, so that this cueing can be avoided.

Conditioning/Training trials

Before determining the MRL, it is important to establish that the infant will respond to supra-threshold stimuli with the appropriate head-turn response.  A clear head-turn, permitting the infant to see the reinforcement object, is the required response. To be considered a response, the head-turn must occur within four seconds of the stimulus presentation. A brief eye-shift in the direction of the reinforcement, or simply looking up at the distracter, is not to be considered a response. The visual reinforcement should be of sufficient duration for the infant to see it briefly (0.5 to 1.0 second); longer reinforcement may lead to extinguishing of the response, especially with infants who are approaching the upper age limit (24 months) of the target population (Culpepper and Thompson, 1994).

Two consecutive, reinforced responses at a supra-threshold level are required to establish that the infant has been conditioned.  It is recommended that the first air-conducted stimulus be presented through the insert earphone in the better ear (if known) at 50 dB HL, or at 10-20 dB above the previously-established ABR EHL threshold for that ear and frequency (if known). The infant may respond spontaneously the first time s/he hears a sound, or may have to be trained to do so. If the child responds by turning his/her head, the reinforcement should be presented. Once the infant’s attention is returned to midline, the same stimulus should be presented again.  If the infant responds again, reinforcement is provided.  These responses must be recorded on the VRA Worksheet. A response is recorded on the worksheet as “+” in the column which designates the stimulus intensity level which has been presented. Absence of response is recorded as “O”.  If the infant does not respond spontaneously to two presentations of the initial signal, the intensity is increased by 20 dB.  If the infant still does not respond, the next signal should be paired with the presentation of the reinforcement.  If the infant does not turn to the reinforcement when it is activated, the distracter may have to point it out to him/her. This pairing of sound and reinforcement may be done twice, then the signal should be presented alone, to see if the infant would now respond to the sound.  If s/he does, two consecutive presentations resulting in clear responses must be given before the search for the MRL is begun.  If there is no response, a listening check of the equipment should be done.  If there is no equipment problem, the intensity level may be increased, or the stimulus frequency or ear of presentation may be changed, and the conditioning repeated until a response is observed.

It is important to avoid pairing the reinforcement with a sound that is inaudible to the infant.  If no responses to sound have been observed, but if the baby turns when the visual reward is presented, it is recommended that a vibrotactile stimulus be used for training, since this is a signal that the infant, regardless of hearing impairment, can almost certainly perceive.  To provide such a stimulus, place and hold the bone-conduction oscillator in the infant’s hand or against his/her arm or leg, and present a 500 Hz bone-conducted tone of 55 dB HL (or 250 Hz BC at 40 dB HL) through the oscillator. Pair this vibrotactile stimulus with the reinforcement, directing the infant’s attention to the reinforcement toy.  Once the head-turn response to the vibrotactile signal has been established, the bone-conduction oscillator is removed, and the auditory stimuli are reintroduced, at increasingly higher intensities until responses are observed.   As soon as the infant has given two consecutive responses at a particular level, MRL search may begin. If the infant shows no interest in the visual reinforcement, and will not turn to look at the reinforcement toys after 5 to 10 trials, further testing with the VRA technique may not be possible or useful on this occasion.

MRL Search

Begin at 20 dB below the level where the infant has just responded twice during the conditioning trials.  If a clear response is noted at this lower level, reinforce and record the response, and lower the intensity by 20 dB for the next stimulus presentation.   If there is no response, raise the intensity by 10 dB. Threshold is bracketed by using a 20 dB-down, 10 dB-up technique. Minimum Response Level (MRL) is defined as the lowest intensity level that elicits two clear responses (out of three presentations). Where the MRL is greater than 30 dB HL, there must be at least one no-response trial at no more than 10 dB below the level reported as MRL. Normal hearing for this test protocol is defined as 30 dB HL, and testing at a given frequency may be discontinued for the tested ear if the criterion for MRL is met at 30 dB HL. (The audiologist may, at his/her discretion, go beyond, i.e. lower than, this level, as long as this does not compromise obtaining MRLs at other required frequencies.   Similarly, the audiologist may use a 5-dB step size when close to MRL, as long as this does not compromise the test objectives.)  The response, or absence of response, to each trial signal must be recorded as “+” or “O” on the Worksheet.

If, when the test ear is changed and a stimulus is presented in the second ear, the infant turns toward the side of the previous reinforcement, this reinforcer is activated. If the infant turns in the direction of the new signal, the corresponding reinforcer is activated.  Any definite head-turn in response to a signal must be reinforced.  However, the use of the left-side reinforcement for left-ear stimulation and the right-side reinforcement for right-side stimulation will probably maintain the infant’s interest longer than the use of reinforcement on a single side. Therefore, when the test ear is changed and the stimulus is presented, if the infant turns to the side of the previous reinforcement, the examiner could illuminate the reinforcer on the side of the stimulus and the distracter could point out the new reinforcer to the infant.

The MRL for each frequency and ear is recorded on the Summary of MRLs form and on the Audiologic Assessment report form.

Control Trials

In order to have a measure of the reliability of the infant’s head-turn responses, control trials are inserted at regular intervals (following a positive response).  The control trial is a specific interval during which the audiologist determines whether a head-turn response occurs in the absence of auditory stimulation.  The purpose is to determine how likely it is that the infant is truly responding to the test sounds, and not just turning his/her head randomly in the hope of receiving reinforcement.

The protocol for the control trial is the same as that for a stimulus trial, except that no tone is presented, and no reinforcement is provided if the baby turns his/her head. The control trial is carried out when the infant is in the appropriate state for a test stimulus presentation, i.e. when s/he is attentive and facing forward. The audiologist informs the distracter of the beginning of the control trial, and the infant’s behaviour is observed for the next four seconds to determine if a head-turn response occurs.  The audiologist records the presence (+) or absence (O) of a response on the worksheet beside the indication for a control trial (C). When an infant responds during a control trial, the audiologist must insert an additional control trial to reassess response reliability.  The more often the infant responds to the test sounds (for example, a baby who responds to all stimuli down to the criterion “normal” level of 30 dB HL), the more important the control trials are in establishing the response reliability. Even though the absence of response to a sound stimulus may represent a type of control trial, it is required that specific, no-sound control trials be used as indicated on the worksheet.

For the test results to be considered reliable, the infant must demonstrate absence of response in at least 70% of the control trials. The number of control trials in which the child gave no head-turn response is expressed as a percentage of the total number of control trials, and is recorded on the Worksheet as ____% correct control trials.  Without formal control trials to determine the response reliability, the validity of the VRA results is open to question, and the results cannot be used with confidence to plan the individual pattern of communication development services.

# Distortion Product Otoacoustic Emissions (DPOAEs)

## Role of DPOAEs

DPOAEs reflect cochlear function, and their normality suggests that the auditory system up to and including the outer hair cells is functioning normally or at least is not markedly abnormal.  They can be reliably recorded in sleeping newborns, given a quiet acoustical environment.  DPOAE measurement is a recommended component of a complete IHP assessment, and it should be done whenever practically feasible.  However, in the event that AC tonepip ABRs are clear and reproducible at the lowest required stimulus levels at both 0.5 and 2 kHz in both ears, the contribution of both DPOAEs and MEA to the overall Assessment is limited.  In any other circumstance than definitely normal tonepip ABRs, then the contribution of both DPOAE and MEA increases substantially, both in terms of the cross-check principle and for refining the overall description of otologic status.  It is especially important that corroborative measures be sought if there is any uncertainty about the reliability of ABR threshold estimates.

## DPOAE Procedure

The current IHP instrumentation (Biologic NavPro) includes diagnostic OAE capability in its ‘Scout’ option. The required test parameters for diagnostic DPOAE measurements will be provided by the IHP.  The current IHP standard protocol includes replicated DPOAE measurements at nominal frequencies of 1, ~1.5,  2, ~3 and 4 kHz.  The f2/f1 ratio is 1.2, with f1 and f2 levels of 65 and 55 dBSPL.

DPOAE measurement at 1 kHz is attempted in the diagnostic context because of the importance of frequency-specific information about cochlear status at low frequencies, but it is especially vulnerable to the generally higher levels of physiologic and environmental noise at low frequencies.  The inclusion of 1 kHz is under continuing evaluation.  Extension of measurements beyond 4 kHz is of questionable clinical utility, and is vulnerable to error arising from standing wave effects.  Also under evaluation is the use of five measurement nominal frequencies spanning two octaves. The known fine structure of both the puretone audiogram and the DP-gram means that on occasion, an individual DPOAE measurement may give a very low amplitude, despite normal cochlear function.  The multiplicity of test frequencies can inform such situations and guard against over-interpretation of DPOAE ‘abnormalities’, and there may be a case for a few more measurements over the two-octave range required.  A revision of the protocol will be issued, if appropriate.

Regardless of the overall DPOAE test outcome, immediate repetition of the test is recommended, to confirm reliability of measurements.  Repetition may be omitted if the DPOAE amplitudes exceed 5 dB and the signal to noise differences exceed 10 dB at 1-4 kHz.  DPOAE testing additional to the IHP protocol is not recommended, because it takes time and adds little relevant information.

## DPOAE Detection

At each nominal DPOAE test frequency, the initial decision after measurement is whether the observed value of DPOAE level represents a genuine DPOAE or is in fact due to noise alone.  It is common to consider both the apparent DPOAE amplitude and the distance from the noise floor in assessing whether an OAE is genuine.  The absolute value of the noise floor is also relevant in determining whether the measurement conditions were such that a normal OAE would be detectable.

Typical criteria for defining a DPOAE to be present, for single stimulus frequency pair, are that its amplitude should exceed –10 dB SPL, and its distance from the noise floor should exceed 5 dB.  However, there is a substantial range of such detectability criteria, as well as considerable variability in both test parameters (such as the amount of averaging) and definitions of measures (such as the signal to noise ratio). There are insufficient data for the IHP to define specific, quantitative criteria.  Furthermore, the detection criteria are inherently statistical, with the usual associated concepts of false-positive and false-negative detection error.  In a rational and quantitative approach to the definition of DPOAE detectability criteria, the costs associated with detection errors would be considered and would affect the criteria chosen.  It is plausible that different criteria of DPOAE detectability should be used in different diagnostic situations, but there is currently little quantitative basis for specific values.  See Brown et al (2000) for illustrative data and discussion.

A preliminary recommendation by the IHP is that a conservative approach be taken generally at individual frequencies, such that a distance of at least 8 dB from the noise floor, an absolute amplitude of at least -10 dB and a test-retest maximum difference of 5 dB be required in order to consider a DPOAE to be definitely present.  Published data suggest that an 8 dB difference criterion will yield a false-positive emission detection rate of about 1%, whereas a 3 dB criterion will give about 10% false positive detection.

Normative noise floor levels have typical 99th percentile values in normal young adults when tested in a soundroom of about –8 , -17 and -21dB at 1, 2 and 4 kHz, respectively.    Observed noise levels much greater than these limit the opportunity for an OAE to be detected reliably.  Noise levels are commonly elevated by 10 dB or more in environments other than audiometric soundrooms, at frequencies below about 2 kHz.

## Interpretation

The display of DPOAE results includes, in graphical and tabular form, the stimulus and OAE amplitudes in real-ear dBSPL, noise floor values in dBSPL and OAE/noise floor differences in dB, for various frequencies of stimulation.  The display also shows the 90th and 95th percentiles of the distribution of amplitude for a population with impaired hearing, and the 5th and 10th percentiles of amplitude for a population with normal hearing.  These are based on large-sample normative data obtained by Gorga et al (1997).  It is clear from the percentile values that the ranges of DPOAE amplitude for the ‘normal’ and ‘abnormal’ populations are substantial, and that the tails of the two distributions overlap considerably.

It follows from the statistical distribution of DPOAE amplitude over subjects that there is a range of DPOAE amplitude that is neither clearly normal nor clearly abnormal.  For the Biologic Scout, the fifth percentiles of the amplitude distributions for young normal adults are in the range about 4-8 dB SPL.  Therefore, any reproducible amplitude in the range about –10 to 5 dB may reflect a genuine DPOAE but with reduced amplitude.

Many factors other than the target PCHI may cause reduction or absence of DPOAES, and some subjects who have the target disorder at some specific frequencies may manifest normal DPOAEs at all frequencies, not just those remote from the target disorder.  Given a quiet subject and an adequate acoustical environment, OAE recording is adversely affected by inappropriate probe placement (such as against the meatal wall), probe blockage (by vernix, cerumen or meatal debris), and the status of the middle ear.  Active middle ear infection, negative middle ear pressure, fluid or debris in the middle ear, and ossicular abnormalities are likely to reduce or abolish the OAE.  In the presence of otitis media, OAEs are rarely recorded unless the air-bone gap is less than about 15 dB.  Furthermore, OAE development is generally (but statistically) a more sensitive indicator of cochlear status than are puretone thresholds, so it is possible to have reduced or absent OAEs even though hearing is normal on conventional audiometric criteria and on IHP criteria.

It follows that DPOAE measurements do not yield quantitative estimates of hearing thresholds and they do not definitively categorize individuals as having normal or impaired hearing.  Rather, there is a statistical, predictive relationship between the amplitude of the DPOAE and the severity and type of hearing impairment.  For hearing levels greater than about 40 dB, an OAE is unlikely to be observed at the frequency of the loss, if the etiology is a cochlear disorder other than auditory dys-synchrony.

Because of the statistical relationship of OAEs to hearing sensitivity, and because the perception of sound requires much more than a functioning auditory periphery, a normal OAE as an isolated finding suggests absence of the target disorder but does not guarantee it.  For example, OAEs are routinely normal in the presence of auditory dys-synchrony or retrocochlear disorders, and a wide range of actual, perceptual hearing thresholds may be seen in such disorders.

Within the constraints of their statistical nature, generally the DPOAE results should be consistent with the threshold findings.  An observation of normal DPOAEs and abnormal thresholds should prompt careful review of the data.  The first question is whether there has been error or misinterpretation.  For example, are the DPOAEs unequivocally present?  Was the absence of ABR at lower stimulus levels genuine?  Were the EEG noise conditions satisfactory or is it possible that the ABR was obscured?  Were the numbers of sweeps and replicates sufficient?

If the DPOAE are definitely present and the thresholds are elevated up to about 35 dB EHL, there is not necessarily any conflict.  But if the thresholds are 40 dBEHL or greater, the likelihood of seeing a clear DPOAE is negligible if the disorder has a conventional cochlear origin.  As the threshold increases, the likelihood of auditory dys-synchrony or a retrocochlear disorder increases significantly.  Indeed, most cases of auditory dys-synchrony show an absent ABR at the highest stimulus levels, even with click stimuli.  In that case, the CM recording and otoneurologic click ABR measurements described earlier are likely to clarify the diagnostic picture.

The most probable cause of a finding of absent or depressed DPOAEs in the presence of normal thresholds is a minor middle-ear disorder. and the MEA may shed light on that possibility.  If MEA is normal, the apparent DPOAE/ABR discrepancy should prompt repeat of the DPOAE with careful attention to probe placement and noise levels.   If these variables appear satisfactory, it is important to consider the possibility that the R thresholds may have been underestimated; the records should be reviewed to look for possible false-positive response detection judgements at the lower levels.  Because it is unlikely, but certainly possible, to see a clear response at 30dBEHL in the presence of slight hearing loss sufficient to degrade or abolish DPOAEs, these apparent contradictions of test outcome should be seen primarily as an indicator for careful review of the findings.  If repeat testing is deemed appropriate in the light of critical review, such testing should generally be very focussed in its objectives and in the range of conditions explored.

Middle Ear Analysis (MEA)

Middle ear analysis (MEA) has a secondary role in the context of IHP Assessment.  It is recommended that MEA be done wherever practically feasible.  The value of MEA ranges from negligible to substantial, depending on specific circumstances, some of which are noted below.

MEA includes tympanometry (measurement of otoacoustic immittance or its components) as well as measurement of middle ear muscle reflexes (MEMR).  Tympanometry is a routine component of conventional audiometric assessment, and its rationale and contribution more generally also apply to the situation of IHP Assessment of infants.  However, some changes in indications, procedure and interpretation are necessary, to reflect the target population and operational context of testing.

## Tympanometry Procedure

All the relevant components of MEA in infants can be performed on the diagnostic immittance instrumentation provided by the IHP (currently the GSI Tympstar Version 2).  Test parameters for use in Assessments will be specified by the IHP.

For infants up to and including six months corrected age:

Tympanometry shall be done using a 1kHz probe frequency, with repetition as necessary and feasible, to improve reliability.

The key abnormality criterion is a compensated peak static admittance of <= 0.6 mmho, compensated from the negative tail at -400 daPa.

For infants over six months corrected age:

Tympanometry shall be done using a 226 Hz probe frequency, with repetition as necessary and feasible, to improve reliability.

The key abnormality criterion in the age range 7-12 months is a compensated peak static admittance of 0.1 mmho, compensated from the positive tail at +200 daPa.  From 13-18 months, the criterion is 0.15 mmho.  From 19 months on, the criterion is 0.2 mmho.

There is evidence that tympanometry with a low-frequency probe is insensitive to the presence of middle-ear fluid in infants under about 8 months of age.  It was suggested by Paradise et al (1976) that in newborns and young infants the meatal wall is distensible, and that this may cause an artifactual tympanometric peak and mask reduced TM compliance due to a middle ear condition.  That hypothesis has been disputed, in favour of a more complex dynamical mechanism of admittance peak generation.  Whatever the mechanism of falsely normal tympanograms, it appears that tympanometry with a high-frequency probe is more sensitive to the presence of middle-ear fluid.  All tympanometric records shall be printed out and retained.

# Middle Ear Muscle Reflexes Procedure

# The measurement of MEMRs is recommended wherever feasible.  The presence or absence of MEMRs shall be measured in the ipsilateral mode with a 1 kHz stimulus and a 678 Hz probe.  There is no age limit on the use of the high-frequency probe.  There is substantial evidence that the likelihood of obtaining a reflex when middle-ear conditions are within normal limits is increased for ipsilateral stimulation and with the use of high-frequency probes.  With contralateral measurements, and with low-frequency probes, reflexes are absent in a high proportion of newborns and young infants with no evidence of a middle-ear disorder.  Therefore, such measurements have little clinical utility, either with respect to middle-ear status or to rule out severe/profound hearing impairment.

# While MEMR measurement at other stimulus frequencies is discretional, the IHP protocol is based on the measurement at 1 kHz.  There is little evidence of incremental diagnostic contribution by testing at several frequencies.  Indeed, there is a risk of increased error because the likelihood of reflex absence unrelated to middle ear conditions increases at higher frequencies and with more testing

Therefore, acoustic reflexes shall be elicited with a 1 kHz stimulus and measured ipsilaterally, using a 1 kHz probe frequency irrespective of age.

Stimulus level shall start at 90 dB and increase in 5 dB steps up to no greater than 100 dB.  Note that for a given nominal level, real-ear SPLs in young infants may be up to 20 dB greater than in adults.

Reflex presence is defined by a clear, negative deflection that is repeatable at any stimulus level.

MEMR Interpretation

# In the context of the IHP Assessment, the clinical utility of the MEMR is primarily to lend support to the ABR measurements.  Misinterpretations of ABR records resulting in drastic overestimation of hearing thresholds have been reported, and while it is anticipated that the expertise, training and protocol within the IHP will reduce the likelihood of such events, every additional precaution within reason should be taken.

# When the results indicate at least a severe hearing impairment, MEMR measurement should be attempted unless there is a contraindication.  In this situation, reliable reflex presence is a significant finding that should result in a critical review of the threshold estimates.  Key issues are whether the measurement conditions were appropriate, particularly with respect to EEG noise levels, the size of averages and the number of replicate averages.  Reflex absence lends weak support to any ABR-based inference of severe or greater sensorineural impairment, except when there is evidence of a conductive component.

# Conventionally, the absence of an MEMR lends support to an inference of a conductive component, provided that absolute AC thresholds are below about 70 dBEHL.  The MEMR is also absent in the presence of auditory dys-synchrony.

# Overall Diagnostic Inference

# The suite of procedures in the core protocol offers many possibilities for evaluation of consistency or discrepancy among measures.  The crucial decision is usually related to confirming the validity and accuracy of the ABR threshold estimates.  The internal validity and reliability of threshold measurements is increased because of the IHP requirement for threshold estimates at 0.5 kHz and 2 kHz at a minimum, the use of both AC and BC routes, and the contingent click ABR measurements in the event of no response at maximum levels.  However, an important principle is that corroborative evidence for the ABR findings should be sought wherever possible.

# The downside in seeking corroborative evidence is that when the corroborative measures such as DPOAE and MEA are themselves error-prone and subject to a host of variables, they must not be allowed to undermine or compromise inappropriately the primary inferences from the testing.  Rather, the corroborative measures provide an indication for critical re-evaluation and where necessary, confirmation of findings.  A balance is always required, and that is the nature of clinical judgement required from IHP audiologists.

# The most probable scenarios that may lead to delay in definitive Assessment are: (i) inability to obtain sufficiently reliable threshold estimates without resorting to sedation or general anesthesia, (ii) audiometric uncertainty arising from evidence of ABR waveform abnormality suggesting auditory dys-synchrony or other disorders that degrade neuronal synchrony in the auditory brainstem, and (iii) audiometric uncertainty due to a transient or fluctuating conductive overlay on a genuine sensory impairment.

# Provided that auditory dys-synchrony and retrocochlear disorders are absent, it is usually possible to obtain reliable threshold estimates with tonepip ABR in ANY infant, provided there is a willingness to consider testing under sedation.  By far the most common cause of inadequate ABR results is poor EEG conditions due to electromyogenic artifacts associated with tension or gross movement.

# When the array of test outcomes suggest the presence of auditory dys-synchrony or a retrocochlear disorder, ABR thresholds based on wave V-V’ cannot be relied upon.  They will tend to overestimate the true perceptual thresholds by an amount that depends on the level of dys-synchrony present.  When neurogenic early waves are present but wave V-V’ is not, the early waves may give some indication of hearing thresholds, but wave I, for example, is limited in its sensitivity and the extent to which the disorder affecting the later ABR waves may also compromise perceptual sensitivity is unknown.

# For infants with Auditory Dys-synchrony or retrocochlear disorders, currently behavioural assessment is usually required to obtain valid estimates of hearing sensitivity.  Reliable thresholds may not be available until VRA becomes accurate and specific.  While there are reports that at least some infants with auditory dys-synchrony may do well with amplification or with cochlear implants, there is at present insufficient information to recommend a specific approach to communication development.  The approach following a determination of auditory dys-synchrony is at the discretion of the IHP audiologist.  The IHP is currently evaluating more advanced assessment methods to address these circumstances.

# Usually, as the infant gets older it will be increasingly important to integrate VRA results into the overall audiometric picture.  During the course of that transition it should not be forgotten that a threshold estimate obtained by VRA represents not perception of sound, but responsiveness to it.  So the first question that arises if VRA results are worse than ABR results obtained with a high-quality protocol is whether there is a responsiveness issue.  The second question is whether the hearing impairment has progressed.  Risk indicators should be reviewed carefully in relation to possible progressive impairment.  Conversely, if the VRA results are markedly better than those from the ABR, the first question is whether there was adequate control of false-positive response in the VRA. A repeat ABR test under sedation may be required to resolve these situations definitively.

# Equally, it should not be forgotten that the ABR is a proxy for perception, and that several factors can compromise the validity of that relationship. The most obvious of these is the possibility of inadequate procedure or interpretive error in the ABR.  There is a need to take the strengths and weaknesses of both behavioural and electrophysiologic assessment into account, to review findings critically, to repeat measures where necessary and finally to take family observations into account, in the pursuit of an accurate overall assessment.

The medical community plays a key role in establishing etiologic diagnosis of the hearing impairment, in assessing treatment needs, in providing medical and surgical treatment of remediable conditions, and in the broader assessment of underlying disorders and conditions that are related directly or indirectly to the hearing impairment.  These processes may include a variety of investigations (such as ophthalmologic, radiologic, metabolic,  neurodevelopmental, genetic) and referrals, such as for social supports.

The IHP is collaborating with the medical community to develop guidelines for medical management of infants identified by the program, and to facilitate effective and efficient means of referral to appropriate physicians and medical services, throughout the province.  It is particularly important to provide family physicians, pediatricians and otolaryngologists with the information they need so that they will support and encourage families to follow program recommendations.  It is also essential to establish fast-track referral mechanisms to otolaryngologists with pediatric experience, and efforts to facilitate these mechanisms have been initiated.

# Confirmed ‘Normal Hearing’

From the IHP perspective, hearing is ‘normal’ when the target disorder is deemed not to be present.  This is not the same thing as the conventional, clinical meaning of ‘normal hearing’.  Clear and reproducible ABRs by air conduction at 0.5 kHz and 2kHz in each ear, at the required minimum levels (currently 25 dBEHL) are sufficient to define ‘normal’ hearing from the IHP perspective.  It is recommended, but not required, that if DPOAEs are clearly abnormal at 3 or 4 kHz, but clearly normal at lower frequencies, then ABR testing at 4 kHz should be considered.  Because there are many causes of absent or depressed DPOAEs, normality of OAEs at all frequencies is not necessary for an overall conclusion of normal hearing in the IHP sense.

When a ‘normal hearing’ determination is made, the family should be counseled fully about what is meant by such a result and about the need for continued vigilance.  The family should be provided with the standard IHP documentary information covering issues such as risk indicators, communication development milestones and what to do if a concern develops.  This information should be provided in the most relevant language currently available from the IHP.

Confirmed 'Permanent' Childhood Hearing Impairment

## Conductive Impairment

Conductive impairment associated exclusively with otitis media does NOT constitute the target PCHI, even if it is chronic.  However, conductive impairment associated with structural disorders such as canal atresia is included in the definition of PCHI. Scheduling of Assessment for infants with colds or symptoms of an active middle ear infection should be deferred wherever feasible.  However, at the Assessment many infants may yield a pattern that suggests a minor, conductive impairment and a probable middle ear disorder.  An example is 0.5 kHz AC ABR thresholds elevated to 40 dBEHL with normal BC thresholds, normal 2 kHz AC ABR thresholds, absent or markedly depressed OAE, questionable or abnormal tympanometry and absent MEMR.  These infants almost certainly do not have the target disorder, and the minor impairment that they do have is likely to be transient.  The required action is at the discretion of the audiologist.  Discharge with a cautionary letter to the primary care physician is an option.

For infants with a more substantive but exclusively conductive hearing impairment, such as with an air-bone gap of 20 dB or more at any frequency in either or both ears, prompt medical referral seems prudent but the IHP has no specific recommendations in this situation.  Local practices and standards of care may take precedence.  In the absence of sensorineural impairment, subsequent audiologic re-assessment related to management of middle-ear disease is outside the scope of IHP funding.

Infants with clearly visible signs of a structural anomaly (such as an atretic ear) should be recommended for referral to an otolaryngologist, if they are not already under such management.  Further Assessment of conductive hearing impairment associated with structural anomalies is within the scope of the IHP.

## Sensorineural or Mixed Impairment

The infant is defined to have the target PCHI by any elevation of BC tonepip ABR threshold of 10 dB or more above the required minimum test levels at any frequency in the range 0.5 through 4 kHz, in either or both ears.  In the event that BC testing has proved unfeasible or inconclusive, AC threshold measurements may serve to define sensorineural hearing levels provisionally, provided that immittance results are clearly normal.  These situations include a wide range of hearing impairments, with diverse implications for communication development options.  Referral to an otolaryngologist with is strongly recommended for ALL infants with significant, permanent hearing impairment, and such referral is normally required prior to prescription of amplification.  The IHP is working to facilitate referral processes, throughout the province.

Some infants will be deemed to have the target PCHI by virtue of a sensorineural component, but will also have a conductive overlay associated with otitis media.  While audiometry directly associated with management of the conductive component technically falls outside the scope of the IHP, the existence of the sensorineural component renders these infants candidates for whatever repeated assessment is necessary to facilitate and expedite the elected communication development options.

Because of the possibility of progression of sensorineural impairment, the importance of audiometric followup irrespective of any other course of events should be fully explained to families at the appropriate time.

On completing the Assessment, the audiologist shall review the results and their implications with the family.  The exchange must reflect the IHP key principles noted earlier.  The information given to families shall address:

- The nature and implications of the Assessment results.

- Identification of IHP family support staff and contact particulars.

- Provision of appropriate IHP documentary information.

- Audiologic aspects of communication development options, including hearing aids, assistive devices, cochlear implants, and communication development strategies.

- Medical referrals.

- Audiologic follow-up.

## Records and Reports

The detailed results of the audiologic Assessment must be maintained in a local clinical chart/file, in full compliance with both current CASLPO standards and IHP specifications.  The records must be fully sufficient to demonstrate compliance with the required elements of the IHP Assessment protocol, given an audit (See audit details for details of required record keeping).  They should also be sufficient to facilitate consultative, clinical review and case conferencing. The records must permit complete specification of the test parameters for all ABR measurements.  Specific requirements for hardcopy and electronic storage of ABR data were outlined earlier.  In addition, the required elements of the results must be entered into the appropriate IHP form and sent to the designated coordinating Preschool Speech and Language agency.

Second Opinions

Because of the many steps (such as special training, standard equipment and protocols, etc) that are being taken to ensure that IHP Assessment services are consistent and of the highest possible quality, second opinions under IHP funding are not encouraged.  It is recognized that in the past, second opinions may have formed part of the process of gaining family acceptance of test results.  However, the program funding envelope is fixed and for every Assessment that is duplicated, another Assessment that may be more crucial will go unfunded.  Audiologists are encouraged to emphasize to families the unprecedented standardization of equipment, procedures and training that forms part of the IHP’s plan to achieve the highest possible service quality.

Families may obtain additional testing outside of the IHP should they wish to do so, but such testing generally will not qualify for IHP funding.  The results of testing by audiologists who are not specifically registered with the IHP, or of testing with non-IHP equipment or protocols, will not normally be considered a valid basis for any further services that may be received within the IHP.  Any such further services will be based on previous procedures and results obtained within the IHP.

Individual families may apply to the IHP for a second opinion, which may include re-interpretation of existing data or a new Assessment.  Application would normally be made through the regional family support worker.  A clear and reasonable argument would be required to justify program funding to duplicate Assessment services.  Consent for program management to review detailed audiologic records, including ABR waveforms, will normally be required.

A step that might assist in resolving some situations would be to offer families who are requesting second opinions the alternative option of special review of the Assessment records, possibly by an expert review panel mediated by IHP management.   The desirability and feasibility of such a process is under evaluation.

Communication Development Options

If a target PCHI is confirmed, the parents/caregivers shall be provided with appropriate information and supports.  They shall be given contact data for IHP family support personnel, who will provide psychological support as well as general information about community supports and options.  The IHP will make available standard, full, evidence-based and unbiased information about ALL available communication development options (See PCHL Guidelines Document for details).  The audiologist is the key informant in relation to amplification, assistive listening devices and cochlear implants, as well as other aspects of communication development.

If the parents/caregivers elect personal amplification, the IHP fully supports its provision before six months of age, wherever possible.   This satisfies the US JCIH’s widely endorsed guideline of completed diagnostic assessment and implementation of an individualized communication development plan by six months of age.

The IHP does not recommend provision of air conduction hearing aids to infants less than three months corrected age, except under special circumstances and at the discretion of the IHP prescribing audiologist.  There are many factors that complicate provision of hearing aids to very young infants.  There is a need for frequent earmold changes and there can be many practical difficulties with fit, feedback and stable attachment. There are also complications with respect to audiometric and prescriptive accuracy, because of rapid changes in ear canal acoustics, in the properties of the middle ear, and in neurodevelopment.  Furthermore, the IHP is unaware of any scientific evidence that hearing aid fitting at, say, one or two months is superior to fitting at four to five months.

Where amplification is indicated and elected by the family, medical referral leading to review by an otolaryngologist is required to establish the absence of medical contraindications to non-medical intervention.  This aspect could be integrated with the requirement identified earlier that referral to an otolaryngologist is desirable wherever feasible, as part of the normal process of overall assessment of an infant determined to have the target PCHI.

# The IHP Amplification Protocol for prescription, fitting, verification and validation of personally-worn amplification, as well as relating to other assistive devices, is described in detail in a separate document.

# Training and Support

The IHP protocol for Assessment requires specialized expertise.  Experience with a protocol comparable to that described here has been obtained in very few centers in Canada and the US.  It follows that current professional experience with key elements of this protocol is limited.  Furthermore, major errors in audiologic assessment have arisen in other UNHS programs internationally.  Such errors have serious consequences for the affected families, as well as for the continuation of UNHS-based programs.  In the IHP context, the likelihood of interpretive error must be minimized, and procedural error must be eliminated.  At the December 2000 workshop, it was unanimously agreed that special training and support mechanisms would be required and would be in the interests of care quality and overall success of the IHP.

Because of the extraordinary importance of high quality and consistent practices in the IHP, a training program has been implemented for audiologists wishing to provide audiologic Assessments for infants under IHP funding.  Each and every individual audiologist wishing to undertake such Assessments must attend for approved training.  The training is provided on an ongoing basis, to develop and maintain the required pool of professional expertise in the Ontario.  Reasonable, direct training expenses (travel, accommodation, materials, etc) are reimbursed by the IHP.

Several mechanisms of ongoing operational support are under consideration.  These include web-based information and educational resources, case-conferencing, internet, fax or phone consultations for interpretation of test results and for other clinical support issues, and additional methods associated with telehealth and distance learning.  Seminars, updates and other traditional avenues for continuing education are also offered.

The clinical support mechanisms will complement other aspects of program quality management.  This will include database review, random audits of clinical records and/or site visits.  These procedures are now widely recognized as appropriate in the quality management of medical and allied health services delivery.

Suggested Readings

Bamford J, Gliddon M, Green R, Munro K, Parry G, Sutton G, Wood S.  Visual Reinforcement Audiometry Testing of Infants: A Recommended Test Protocol. Day J (Ed.). British Neonatal Hearing Screening and Assessment Program; October, 2001.

Brown DK, Tobolski C, Shaw G, Dort J. Towards determining distortion product otoacoustic emission protocols for newborn hearing screening.  J Sp-Lang Pathol Audiol 2000;24:68-73.

Cone-Wesson B, Ramirez G. Hearing sensitivity in newborns estimated from ABRs to bone-conducted sounds.  J Am Acad Audiol 1997;8(5):299-307.

Cone-Wesson B, Vohr BR, Sininger YS, Widen JE, Folsom RC, Gorga MP, Norton SJ. Identification of neonatal hearing impairment: infants with hearing loss. Ear & Hearing 2000;21(5):488-507.

Gorga MP, Neely S, Ohlrich B et al. From laboratory to clinic: a large-scale study of distortion product otoacoustic emissions in ears with normal hearing and ears with hearing loss.  Ear & Hearing 1997;18(6):440-455.

Gravel JS.  Potential pitfalls in the audiological assessment of infants and young children.  In Seewald RC and Gravel JS (Eds.):  A Sound Foundation through Early Amplification 2001: Proceedings of the Second International Conference.  Phonak AG, 2002, pp 85-101.

Hunter LL, Margolis RH.  Multifrequency tympanometry – current clinical application.  Am J Audiol 1992;July:33-43

JCIH 2000.  Joint Committee on Infant Hearing – Year 2000 Position Statement.  Principles & Guidelines for Early Hearing Detection & Intervention Programs.

Kruger, B. An update on the external ear resonances in infants and young children, Ear and Hearing, 1987:8(6), 333-336.

McMillan PM, Bennett M, Marchant C, Shurin P.  Ipsilateral and contralateral acoustic reflexes in neonates.  Ear & Hearing 1985;6(6):320-324.

Moodie, K.S., Seewald, R.C., & Sinclair, S.T. Procedure for predicting real ear hearing aid performance in young children, American Journal of Audiology,1994:3, 23-31.

Nozza RJ. Developmental psychoacoustics: science to practice. In Seewald RC and Gravel JS (Eds.): A Sound Foundation through Early Amplification 2001: Proceedings of the Second International Conference.  Phonak AG, 2002:37-46.

Sininger YS, Abdala C, Cone-Wesson B.  Auditory sensitivity of the human neonate as measured by the auditory brainstem response.  Hearing Res 1997;104:27-38.

Stapells DR. Threshold estimation by the tone-evoked auditory brainstem response: a literature meta-analysis.  J Sp-Lang Pathol Audiol 2000;24:74-83.

Stapells DR. Frequency-specific evoked potential audiometry in infants. In RC Seewald (Ed): A Sound Foundation through Early Amplification. Phonak AG, 2000, pp 13-32.

Stapells DR, Oates P.  Estimation of the pure-tone audiogram by the auditory brainstem response: A review.  Audiology & Neuro Otology 1997;2:257-280.

Widen JE, Folsom RC, Cone-Wesson B, Carty L, Dunnell JJ, Doebsell K, Levi A, Mancl L, Ohlrich B, Trouba S, Gorga MP, Sininger YS, Bohr BR, Norton SJ.  Identification of neonatal hearing impairment: hearing status at 8 to 12 months corrected age using a visual reinforcement audiometry protocol.  Ear & Hearing 2000;21(5):471-487.

Yang EY, Stuart A.  The contribution of the auditory brainstem responses to bone-conducted stimuli in newborn hearing screening.  J Sp-Lang Pathol Audiol 2000;24:84-91.

Appendix A

IHP Risk indicators

Perinatal

The following neonatal indicators a-m are usually associated with attendance in a special care nursery.  ANY ONE of the indicators is sufficient to place the baby at risk, so if an indicator is positive, it is not necessary to continue to search for the other indicators. However, screening personnel will be asked to identify at least one specific indicator in any baby deemed at risk, for subsequent follow-up and data analysis.

a.	Birthweight less than 1200 grams

b.	Five-minute APGAR score less than or equal to 3

c.	Congenital Diaphragmatic Hernia (CDH)

d.	Persistent Pulmonary Hypertention of the Newborn (PPHN)

e.	Hypoxic-Ischemic Encephalopathy (HIE), Sarnat II or III

f.	Intra-ventricular Haemorrhage (IVH), Grade III or IV

g.	Peri-ventricular Leukomalacia (PVL)

- Extra-Corporeal Membrane Oxygenation (ECMO) or inhaled Nitrous Oxide (iNO) or High-Frequency Oscillatory (HFO) or Jet (HFJ) ventilation

- Hyperbilirubinemia >= 400 muM/L OR exchanged

- Serologically proven AND symptomatic Cytomegalovirus (CMV) infection

- Other proven perinatal TORCHES infection (toxoplasmosis, rubella, herpes, syphilis)

- Serologically proven meningitis, irrespective of the pathogen

- Familial Permanent Childhood Hearing Impairment

- Craniofacial anomaly

- Other high risk indicator specified by baby's treating physician

Infant (0-24 months)

All of the above, plus:

- Parent/Caregiver concern about hearing/speech/language.

- Postnatal infections associated with sensorineural hearing loss (e.g. bacterial meningitis).

- Syndromes associated with progressive hearing loss (neurofibromatosis, osteopetrosis, Usher’s, etc).

- Neurodegenerative disorders (e.g. Hunter syndrome) and sensory motor neuropathies (e.g. Friedreich’s ataxia, Charcot-Marie-Tooth syndrome).

- Head trauma sufficient to cause unconsciousness or skull fracture

Newborns known to be at risk on any Perinatal indicator shall be screened only by AABR.

Infants who manifest any indicator 1-12 are targeted for surveillance procedures through the first two years.

Appendix B

ABR Technical Parameters - adapted from Stapells (2000).

# EEG Recording

Vertex or high forehead to each mastoid.  AC: 1 channel.  BC: 2 channels.

Filters: High Pass 30 Hz,  Low Pass 1500 Hz, 6 or 12 dB/octave, analog.

Artifact reject: trials exceeding +-25 muV or +-two standard deviations of the quietest available EEG, whichever is the smaller.

2000 sweeps per average.  Multiple averages, at least two and often three.

Analysis window length: 25 ms.

# Stimuli

2-1-2 cycle linearly gated tones.  Alternating onset polarity.  Rate 37-41 per second.

Ipsilateral masking is not required.  Contralateral masking is optional.

SPL values (ppe) for 0 dBnHL:

Air Conduction

500 Hz	TDH49  25 dB		ER3A  22 dB

1000 Hz	TDH49  23 dB		ER3A  25 dB

2000 Hz	TDH49  26 dB		ER3A  20 dB

1000 Hz	TDH49  29 dB		ER3A  26 dB

Bone Conduction

Radioear B70 oscillator, PPE force values re: 1 dyne RMS, for 0 dBnHL:

500 Hz	to be confirmed

2000 Hz	to be confirmed

4000 Hz	to be confirmed

Appendix C

VRA Technical Summary - Adapted from Judy Widen, et al (2000)

Pulsed, warbled tones of 1-2 seconds’ duration.

Vary inter-stimulus interval (ISI); longer ISI initially if random head-turns are frequent.

BEGIN with 2000 Hz warbled tone in insert phone (or most sensitive frequency in better ear, if known).

50 dB HL… if baby turns naturally, reinforce

2 correct consecutive responses, go to TEST PROTOCOL.

50 dB HL… no head-turn

70 dB HL… if head-turn, reinforce

2 consecutive responses, go to TEST PROTOCOL

… if no head-turn, go to CONDITIONING TRIALS

CONDITIONING TRIALS

70 dB HL paired with reinforcement, 2 times

70 dB HL "probe"…if head-turn, reinforce

… 2 consecutive head-turns prior to reinforcement, go to TEST PROTOCOL

… if no head-turn on probe, do listening check of earphone. If okay, go to

90 dB HL paired with reinforcement, 2 times

90 dB HL "probe"…if head-turn, reinforce

… 2 consecutive responses, go to TEST PROTOCOL

If no turn on probe, hearing problem or conditioning problem?

Change stimulus frequency? Change stimulus type (e.g. speech)? Change ear? Change stimulus modality (bone-conductor held in hand)? Try sound field?

TEST PROTOCOL  … after 2 consecutive head-turns prior to reinforcement

Down 20 dB, up 10 dB for MRL search

Insert Control Trials according to Worksheet schedule

Test down to 30 dB HL (2 responses out of 3 presentations) OR

Test down to lowest level at which 2 responses out of 3 presentations are obtained.

2nd frequency: 500 Hz in same ear

begin at level of previous response

if 2 consecutive responses, continue MRL search

if no response, increase intensity until response obtained 2 times

continue MRL search

Second ear

2000 Hz at 50 dB HL

… if head-turn (either side), reinforce on side of turn

proceed with MRL search

… if no head-turn, increase intensity until response obtained 2 times

continue MRL search

500 Hz - proceed as above for 2000 Hz

3rd and 4th frequencies (3000/4000 Hz and 1000 Hz) - as for 1st and 2nd frequencies, in each ear.

Deviations from this order may be made if infant begins to habituate:

- change stimuli, or re-condition at a level at which infant responded previously

Bone-conduction

For at least one frequency where AC MRL is >30 dB HL bilaterally.

Vibrator on mastoid on the side of the better AC MRL

Start with intensity at or below air-conduction MRL

Use same test protocol to find MRL

Appendix D

Calibration and Protocol Files – NavPro Scout EP

Basic outline of the calibration verification procedure undertaken for the Bio-logic Nav Pro ABR systems.     Presently there are no ANSI or International standards for ABR instruments, so Bio-Logic calibration/service guidelines are used.      Custom stimulus levels are accommodated.

Preliminary Check procedure:

-Ensure computer boots and functions normally.

-Open Nav Pro program, checking for proper initialization of Nav Pro module.

-Check for revision of software (require 7.26 or higher) to utilize some of the built in calibration check functions.

-Perform System loop test procedure with Bio-Logic supplied adaptor cable, checking Impedance measure accuracy. (±10% allowance)

-Check amp gain changes, Linearity of amps, phase of 2 channels with loop test cable.

Check continuity of electrode leads for breaks, poor condition.

Click Stimulus Measurement:

Insert/Phones:   Peak (rms) Hold value measured @19.1 rate (+35dB, ±5 dB allowance). Bio-Logic reference level.

Bone:   Peak (rms) hold value measured @19.1 rate (+35dB, ±5 dB allowance). Bio-Logic reference level.

Tone Burst Measurement:

Insert/Phones:   Measurement of 500 Hz, 1000 Hz, 2000 Hz, 4000 Hz levels measured with IHP custom levels (re ppe rms or Peak rms ±3 dB allowance), using appropriate Rise/fall & Plateau settings for each frequency.

Bone:   Measurement of 500 Hz & 2000 Hz with IHP custom levels (re ppe rms or Peak rms ±3 dB allowance), using appropriate Rise/fall & Plateau settings for each frequency.

THD and attenuator Linearity can be checked if version 7.26G software or higher.

Miscellaneous:

Instruments will be labeled with emi calibration label to show instrument has met criteria listed.

A copy of invoice will be forwarded to IHP for tracking purposes.

Deviations outside the allowable limits will be forwarded with invoice, and actions taken if applicable.

IHP Protocol File nrm.

IHP Audiology Assessment Protocol Addendum 2003

Scout Diagnostic DPOAE Printouts

IHP recommends that in the Audiologic Assessment, DPOAE tests should be repeated asap, immediately if possible, and that the reproducibility or consistency of the two sets of values be considered (along with absolute SPL and SNR) in the overall judgment as to whether OAEs are present and within normal limits, depressed or absent.

It is recommended that the replicated tests be integrated so that a single plot/printout shows double curves for each type of measure, with left and right test plots side by side and two sets of numerics listed below.  At Mount Sinai, we find this by far the easiest data presentation to assess.  In particular, it makes it much easier to evaluate reproducibility, compared to looking at separate printouts on separate pages.

Instructions for Scout OAE Overlap Plots/Printouts

- In the Scout program you need to access the ‘Open Scout Patient File’ window.  Do this either by selecting ‘File’ in the top toolbar followed by ‘Open Patient File’ -or- select the ‘Open’ icon in the second row toolbar which is represented by the yellow open folder.

- Select all the OAE files for a patient that you wish to view on a single printout.  To do this, hold down the CTRL key on your keyboard and then left-click your mouse over each .OAE file you need.  You should now have several .OAE files highlighted in blue.

- In the same window look for a box titled ‘Multiple File Selection View’.  Within that box you must click on the button/circle beside the heading ‘Right/Left Side by Side’.  Now click on Open.

- You will now have a graphical display of all tests selected with the left ear and right ear results side by side, followed by a detailed descriptive of each test.  Click on the printer icon to receive a printout of this display.

Appendix E

| Age (m) | 250 | 500 | 750 | 1000 | 1500 | 2000 | 3000 | 4000 | 6000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 3 | 7 | 9 | 11 | 13 | 14 | 15 | 20 | 22 |
| 2 | 3 | 7 | 9 | 11 | 13 | 14 | 15 | 19 | 21 |
| 3 | 3 | 7 | 9 | 11 | 13 | 13 | 14 | 19 | 21 |
| 4 | 3 | 7 | 9 | 11 | 12 | 13 | 14 | 19 | 20 |
| 5 | 3 | 7 | 9 | 11 | 12 | 13 | 14 | 18 | 20 |
| 6 | 3 | 7 | 9 | 11 | 12 | 13 | 13 | 18 | 19 |
| 7 | 3 | 7 | 9 | 11 | 12 | 12 | 13 | 17 | 19 |
| 8 | 3 | 7 | 9 | 11 | 12 | 12 | 13 | 17 | 18 |
| 9 | 3 | 7 | 9 | 11 | 12 | 12 | 12 | 16 | 18 |
| 10 | 3 | 7 | 9 | 11 | 11 | 12 | 12 | 16 | 17 |
| 11 | 3 | 7 | 9 | 11 | 11 | 12 | 12 | 16 | 17 |
| 12 | 3 | 7 | 9 | 11 | 11 | 11 | 11 | 15 | 16 |
| 13 | 3 | 7 | 9 | 11 | 11 | 11 | 11 | 15 | 16 |
| 14 | 3 | 7 | 9 | 11 | 11 | 11 | 11 | 14 | 15 |
| 15 | 3 | 7 | 9 | 11 | 10 | 11 | 10 | 14 | 15 |
| 16 | 3 | 7 | 9 | 11 | 10 | 10 | 10 | 14 | 14 |
| 17 | 3 | 7 | 9 | 11 | 10 | 10 | 10 | 13 | 14 |
| 18 | 3 | 7 | 9 | 11 | 10 | 10 | 9 | 13 | 13 |
| 19 | 3 | 7 | 9 | 11 | 10 | 10 | 9 | 12 | 13 |
| 20 | 3 | 7 | 9 | 11 | 9 | 9 | 9 | 12 | 12 |
| 21 | 3 | 7 | 9 | 11 | 9 | 9 | 8 | 11 | 12 |
| 22 | 3 | 7 | 9 | 11 | 9 | 9 | 8 | 11 | 11 |
| 23 | 3 | 7 | 9 | 11 | 9 | 9 | 8 | 11 | 11 |
| 24 | 3 | 7 | 9 | 11 | 9 | 9 | 9 | 11 | 11 |
| 25 | 3 | 7 | 9 | 11 | 10 | 10 | 9 | 12 | 13 |
| 26 | 3 | 7 | 9 | 11 | 10 | 10 | 9 | 12 | 13 |
| 27 | 3 | 7 | 9 | 11 | 10 | 10 | 9 | 12 | 13 |
| 28 | 3 | 7 | 9 | 11 | 10 | 10 | 9 | 12 | 13 |
| 29 | 3 | 7 | 9 | 11 | 10 | 10 | 9 | 12 | 13 |
| 30 | 3 | 7 | 9 | 11 | 10 | 10 | 9 | 12 | 13 |
| 31 | 3 | 7 | 9 | 11 | 10 | 10 | 9 | 12 | 13 |
| 32 | 3 | 7 | 9 | 11 | 10 | 10 | 9 | 12 | 13 |
| 33 | 3 | 7 | 9 | 11 | 10 | 10 | 9 | 12 | 13 |
| 34 | 3 | 7 | 9 | 11 | 10 | 10 | 9 | 12 | 13 |
| 35 | 3 | 7 | 9 | 11 | 10 | 10 | 9 | 12 | 13 |
| 36 (3 yrs) | 3 | 7 | 9 | 11 | 10 | 10 | 9 | 12 | 13 |
| 42 (3.5 yrs) | 3 | 7 | 9 | 11 | 10 | 10 | 9 | 12 | 13 |
| 48 (4 yrs) | 3 | 7 | 9 | 11 | 10 | 10 | 9 | 12 | 13 |
| 54 (4.5 yrs) | 3 | 7 | 9 | 11 | 10 | 10 | 9 | 12 | 13 |
| 60 (5 yrs) | 3 | 7 | 9 | 11 | 10 | 10 | 9 | 12 | 13 |

| Age (m) | 250 | 500 | 750 | 1000 | 1500 | 2000 | 3000 | 4000 | 6000 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 6 | 10 | 12 | 15 | 17 | 16 | 17 | 17 | 20 |
| 2 | 6 | 10 | 12 | 15 | 16 | 16 | 17 | 17 | 20 |
| 3 | 6 | 10 | 12 | 15 | 16 | 16 | 17 | 17 | 20 |
| 4 | 6 | 10 | 12 | 15 | 16 | 15 | 17 | 16 | 19 |
| 5 | 6 | 10 | 12 | 14 | 16 | 15 | 16 | 16 | 19 |
| 6 | 6 | 10 | 12 | 14 | 16 | 15 | 16 | 16 | 19 |
| 7 | 6 | 10 | 12 | 14 | 16 | 15 | 16 | 16 | 18 |
| 8 | 6 | 10 | 12 | 14 | 16 | 15 | 16 | 16 | 18 |
| 9 | 6 | 10 | 12 | 14 | 16 | 15 | 15 | 15 | 18 |
| 10 | 6 | 10 | 12 | 14 | 15 | 15 | 15 | 15 | 18 |
| 11 | 6 | 10 | 12 | 14 | 15 | 15 | 15 | 15 | 17 |
| 12 | 6 | 10 | 12 | 14 | 15 | 15 | 15 | 15 | 17 |
| 13 | 6 | 10 | 12 | 14 | 15 | 15 | 14 | 15 | 17 |
| 14 | 6 | 10 | 12 | 14 | 15 | 14 | 14 | 14 | 17 |
| 15 | 6 | 10 | 12 | 14 | 15 | 14 | 14 | 14 | 16 |
| 16 | 6 | 10 | 12 | 14 | 15 | 14 | 14 | 14 | 16 |
| 17 | 6 | 10 | 12 | 14 | 15 | 14 | 14 | 14 | 16 |
| 18 | 6 | 10 | 12 | 14 | 15 | 14 | 13 | 14 | 16 |
| 19 | 6 | 10 | 12 | 13 | 14 | 14 | 13 | 13 | 15 |
| 20 | 6 | 10 | 12 | 13 | 14 | 14 | 13 | 13 | 15 |
| 21 | 6 | 10 | 12 | 13 | 14 | 14 | 13 | 13 | 15 |
| 22 | 6 | 10 | 12 | 13 | 14 | 14 | 12 | 13 | 14 |
| 23 | 6 | 10 | 12 | 13 | 14 | 14 | 12 | 13 | 14 |
| 24 | 6 | 10 | 12 | 13 | 14 | 14 | 12 | 12 | 14 |
| 25 | 6 | 10 | 12 | 13 | 14 | 13 | 12 | 12 | 14 |
| 26 | 6 | 10 | 12 | 13 | 14 | 13 | 11 | 12 | 13 |
| 27 | 6 | 10 | 12 | 13 | 13 | 13 | 11 | 12 | 13 |
| 28 | 6 | 10 | 12 | 13 | 13 | 13 | 11 | 12 | 13 |
| 29 | 6 | 10 | 12 | 13 | 13 | 13 | 11 | 11 | 13 |
| 30 | 6 | 10 | 12 | 13 | 13 | 13 | 10 | 11 | 12 |
| 31 | 6 | 10 | 12 | 13 | 13 | 13 | 10 | 11 | 12 |
| 32 | 6 | 10 | 12 | 12 | 13 | 13 | 10 | 11 | 12 |
| 33 | 6 | 10 | 12 | 12 | 13 | 13 | 10 | 10 | 12 |
| 34 | 6 | 10 | 12 | 12 | 13 | 13 | 9 | 10 | 11 |
| 35 | 6 | 10 | 12 | 12 | 13 | 12 | 9 | 10 | 11 |
| 36 (3 yrs) | 6 | 10 | 12 | 12 | 13 | 12 | 9 | 10 | 11 |

APPENDIX F

# RECD Measurement Using the Audioscan RM500 (Actual measured)

System Setup: (assuming the system is on and both microphones are calibrated).

1.  	In REM mode, press <ADVANCED FEATURES>

2. 	In the Tests menu select Speechmap/DSL, press <CONTINUE>

3. 	Press <AGRAM>.  When the menu appears, set the parameters by using the arrow and cursor keys. Select the appropriate age (in months) of your patient and indicate the transducer used to collect hearing thresholds (i.e., insert+foam).  Select UCL - avg, and RECD - measure.  Press <CONTINUE>

- The next screen requires that you enter an audiogram.  Enter your patient’s audiogram

and press <CONTINUE>

- Plug the RECD transducer (or EAR-3A insert phone) into the auxiliary speaker jack (it’s usually red and behind the printer).

Coupler Measure: (TIP: These instructions will appear on the screen.)

1.  	Carefully plug the coupler microphone into the jack inside the test box.

2.  	Screw the HA-2 coupler onto the coupler microphone.

3.  	Couple the tip of the insert phone (or RECD transducer) to the tubing of the HA-2 coupler.

4.  	Press <CONTINUE> to introduce the stimulus.

5.  	Press <CONTINUE> again to store the coupler measure. (TIP: once you have done this, it will be saved until the machine is shut off.  It saves time to do this in the morning right after calibrating the two microphones, so that it is completed when the infant arrives.)

Real-Ear Measure:

1.  	Unplug the coupler microphone and plug the probe microphone into the jack inside the test box.

2.	Perform an otoscopic examination on the infant.

3.  	Place the real ear microphone module over the infant’s ear and adjust for length.

4.  	To ensure that the probe tube is positioned correctly, tape a probe tube alongside the foam tip or probe tip so that the tube protrudes about 2mm past the end of the tip. Connect the RECD transducer the tip/tube combination.  Insert the combination into the infant’s ear.  This is helpful in coordinating insertion and ensuring 2 mm remains at the edge probe tip.

5.  	Press <CONTINUE> to introduce the signal.  As it is ‘sweeping’, you will see four curves:

a)  Top curve: the real ear response that is sweeping as you measure the RECD

b)  2nd curve: the coupler response

c)  3rd curve: the difference between a and b (this is the RECD)

d)  4th curve: a dotted line represents an average adult RECD for comparison

6.	Press <PRINT SCREEN> to obtain the RECD values at each frequency.

# Entering predicted RECDs into the Audioscan RM500  (move to amplification)

System Setup: (assuming that the system is on, and that both microphones are calibrated).

1.  	In REM mode, press <ADVANCED FEATURES>

2. 	In the Tests menu select Speechmap/DSL, press <CONTINUE>

3. 	Press <AGRAM>.  When the menu appears, set the parameters by using the arrow and cursor keys. Select the appropriate age (in months) of your patient and indicate the transducer used to collect hearing thresholds (i.e., insert+foam).  Select UCL - avg, and RECD - enter.  Enter age-appropriate normative data from norms provided in appendix E.  Press <CONTINUE>

4. 	The next screen requires that you enter an audiogram.  Enter your patient’s audiogram and press <CONTINUE>

5.	Plug the RECD transducer (or EAR-3A insert phone) into the auxiliary speaker jack (it’s usually red and behind the printer).

Coupler Measure: (TIP: These instructions will appear on the screen.)

1.  	Carefully plug the coupler microphone into the jack inside the test box.

2.  	Screw the HA-2 coupler onto the coupler microphone.

3.  	Couple the tip of the insert phone (or RECD transducer) to the tubing of the HA-2 coupler.

4.  	Press <CONTINUE> to introduce the stimulus.

5.  	Press <CONTINUE> again to store the coupler measure. (TIP: once you have done this, it will be saved until the machine is shut off.  It saves time to do this in the morning right after calibrating the two microphones, so that it is completed when the infant arrives.)

Real-Ear Measure:

1.  	Unplug the coupler microphone and plug the probe microphone into the jack inside the test box.

2.	Perform an otoscopic examination on the infant.

3.  	Place the real ear microphone module over the infant’s ear and adjust for length.

4.  	Insert the probe tube into the infant’s ear using a probe tube insertion depth of approximately 5 mm.   To ensure that the probe tube lies flat on the floor of the ear canal couple a foam tip to the insert phone and RECD transducer using hearing aid moisture-guard wrap.  Insert the foam tip and probe tube combination into the infant’s ear.  This is helpful in coordinating insertion and ensuring 2 mm remains at the edge probe tip.

6.  	Press <CONTINUE> to introduce the signal.  As it is ‘sweeping’, you will see four curves:

a)  Top curve: the real ear response that is sweeping as you measure the RECD

b)  2nd curve: the coupler response

c)  3rd curve: the difference between a and b (this is the RECD)

d)  4th curve: a dotted line represents an average adult RECD for comparison

7.	Press <PRINT SCREEN> to obtain the RECD values at each frequency.

RECD Measurement Using the Verifit

1.	Select <TESTS>

2.	In Test selection menu select <Speechmap>:

Mode:                 REM

Presentation: user preference

Format: Table

Scale (dB): HL

3. 	In <Audiometry> menu, select:

Targets:  DSL

Age:	months

Transducer:  insert + foam

UCL:	AVG

RECD:   Measure

4.	Follow prompt to enter audiometry and measured RECD values

Measure coupler by following prompt directions in test box

Measure RECD by coupling probe tube to probe tip during ABR measurement

5.	Press <PRINT> to obtain a record of the infant’s RECD.

# Entering predicted RECDs into the Audioscan Verifit  (move to amplification)

System Setup: (assuming that the system is on, and that both microphones are calibrated).

1.	Select <TESTS>

2.	In Test selection menu select <Speechmap>:

Mode:  S-REM

Presentation: user preference

Format: Table

Scale (dB): HL

3. 	In <Audiometry> menu, select:

Targets:  DSL

Age:	months

Transducer:  insert + foam

UCL:	AVG

RECD:   Enter

4.	Follow prompt to enter thresholds and normative RECD for age of the infant by month	 (see appendix E for values)

IHP Assessment protocol Draft Dec 17 2004

# Text boxes

RECD Predictions for Foam Tips

11/25/2003

Frequency (Hz)

RECD Predictions for Earmolds

11/25/2003

Frequency (Hz)